## Instructions
DAPHNE DCS intake | Full board, service, network, I2C/SPI, power, thermal, optical, timing, AFE, channel, readout, authority, and command registry
Canonical candidate namespace | urn:dune:pds:daphne
Instance rule | Instantiate {BoardId} and child index placeholders from HWDB/deployment records; do not hand-maintain 192+ board copies.
Contract status | Every row states whether it is current OPC-UA, current Protobuf, locally readable, required production work, or excluded from routine DCS.
Chosen ownership architecture | DAQ and SC are authenticated clients of one neutral PDS OPC-UA gateway. The gateway alone maintains the DAPHNE Protobuf/ZMQ backend, serializes board mutations, and enforces the Control Owner policy; SC never receives or executes a DAQ configuration handoff.
Ownership rule | DAQ owns timing, AFE/channel operating values, triggers, spy buffers, packetization and readout. SC owns power/bias enable, rails, services, firmware lifecycle, network, fans, equipment health and safe state. Column AB is normative.
Write rule | The OPC-UA gateway authorizes every typed method against the exported workbook policy and serializes accepted board mutations; the board service retains independent range/safety validation. All writes require authenticated authority, state/permit guards, bounded execution, audit, and verified readback/postcondition.
Threshold policy | No numeric alarm/interlock thresholds are guessed. PDS/SC/DPS must fill reviewed limits and actions before approval.
External compatibility evidence | The NP02 Compatibility sheet transcribes a supplied 2025 color matrix as historical admission evidence. Red means an activation pair is incompatible; green does not bypass other guards; orange 'probably compatible' remains unresolved and is not a permit. The source is not a production cause-and-effect specification.
Prepared | 2026-08-17
Document status | Draft 0.7 for joint PDS / Slow Controls / DAQ / DPS review
Companion document | DUNE Interface Control Document: DAQ & Photon Detector System (PDS), proposed v8.0 revision
Current prototype namespace | urn:dune:pds:np04; the production namespace and endpoint assignments remain approval items.
EDMS template basis | DCS_Intake_Template.xlsx; all example values and actions are explicitly cleared.
How to filter | Filter Tag List by Subsystem, DCS Access and Control Owner. The OPC-UA Control Policy sheet and companion CSV are the machine-readable authorization input for the gateway. Raw Register Map is evidence, not a request to expose unrestricted MMIO.
Source baseline | Ownership was checked against EDMS 2580777, EDMS 2395364 v2, EDMS 2209895 v1, upstream daphnemodules develop edf75a0, and the reviewed DAPHNE firmware working interface. The NP02 matrix is separately identified as historical evidence; see Source Audit.
Column reference
Column | Meaning
Signal / Tag Name | Short machine name for the tag.
Description | Plain-language description of what the signal is.
OPC-UA Node Address | Full OPC-UA NodeId or browse path the DCS will read.
Data Type | OPC-UA data type. Use the dropdown.
Eng Units | Engineering units ONLY (degC, V, A, psia, bar...). Do NOT put units in the value columns.
Signal Source | Direct = read straight from device. Derived = computed.
Scaling / Derivation | Raw->scaling (slope/offset) OR the formula for derived signals. e.g. (v0-v1)/0.001 ; or raw*0.1-40.
Sensor Span Min | Instrument range low (full measurable span). Used for scaling. Numeric only.
Sensor Span Max | Instrument range high (full measurable span). Numeric only.
Nominal Value | Expected value in normal operation. Numeric only (units go in Eng Units).
Desired Range Min | Low end of normal operating band. Numeric only.
Desired Range Max | High end of normal operating band. Numeric only.
Low Interlock (Trip) | Trip/safety action point on the low side. Leave blank if none. Numeric only.
Low-Low | Second-stage low alarm. Blank if none. Numeric only.
Low | First-stage low alarm. Blank if none. Numeric only.
High | First-stage high alarm. Blank if none. Numeric only.
High-High | Second-stage high alarm. Blank if none. Numeric only.
High Interlock (Trip) | Trip/safety action point on the high side. Blank if none. Numeric only.
Alarm Deadband / Hysteresis | How far the value must recover before the alarm clears. Numeric only.
Alarm On-Delay (s) | Seconds the condition must persist before the alarm fires.
Interlock Action | If this signal trips an interlock, what action must occur?
Access | Read-Only = DCS monitors only. Read-Write = DCS can also set it. Write-Only = command only. Use the dropdown.
Historize? | Should the DCS log this to the historian? Yes/No.
Min Historization Rate | Slowest acceptable logging interval for the historian. Enter an interval in ms, or 'On-Change'. Independent of poll rate.
Poll / Scan Rate (ms) | How often the DCS should read this signal from the device, in milliseconds. e.g. 1000. This is the device poll rate, not the log rate.

## Coverage Summary
Metric | Count
CURRENT OPC-UA POC | 14
CURRENT PROTOBUF | 105
LOCAL READABLE | 44
REQUIRED PRODUCTION | 214
EXPERT/DAQ ONLY | 0
EXCLUDED FROM ROUTINE DCS | 6
Read-only patterns | 362
Write-only method patterns | 21
Control owner: DAQ | 99
Control owner: DAQ / DTS | 27
Control owner: Slow Controls | 195
Control owner: Slow Controls / DPS | 10
Control owner: PDS expert (maintenance only) | 2
Control owner: Board service (shared audit) | 15
Control owner: HWDB / network authority; SC monitors | 17
Control owner: PDS OPC-UA gateway service | 18
Total tag patterns | 383
Raw register patterns | 43
Protocol families | 19
Operations constraints | 10
NP02 compatibility pairs | 15
NP02: Incompatible | 6
NP02: Compatible | 5
NP02: Probably compatible | 4
Interpretation | Counts are variable patterns, not installed tag instances. {BoardId}/{Afe}/{Channel}/other placeholders are expanded from controlled inventories.
Approval blocker | Alarm/interlock numeric limits, production certificates/identities, final endpoints, complete gateware-health block, monitoring cadence, and the detailed DAQ-SC-DPS permit/failure state machine require joint sign-off. The ownership split is fixed by this revision.

## Tag List
Subsystem | Signal / Tag Name | Description | OPC-UA Node Address | Data Type | Eng Units | Signal Source | Scaling / Derivation | Sensor Span Min | Sensor Span Max | Nominal Value | Desired Range Min | Desired Range Max | Low Interlock (Trip) | Low-Low | Low | High | High-High | High Interlock (Trip) | Alarm Deadband / Hysteresis | Alarm On-Delay (s) | Interlock Action | DCS Access | Historize? | Min Historization Rate (ms) | Poll / Scan Rate (ms) | Notes | Control Owner
Endpoint Status | Connected | Bridge can reach the board ZMQ service. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Connected | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=NP04 OPC-UA bridge 0.3 / cached DAPHNE poll ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Stale data ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | StatusSuccess | Most recent board status transaction succeeded. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Success | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=NP04 OPC-UA bridge 0.3 / cached DAPHNE poll ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | StatusMessage | Bounded diagnostic from the most recent status transaction. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Message | String | N/A | Direct | None | [14 empty] | Read-Only | No | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=NP04 OPC-UA bridge 0.3 / cached DAPHNE poll ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Status ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | LastUpdate | Source time of the most recent completed board status update. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.LastUpdate | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=NP04 OPC-UA bridge 0.3 / cached DAPHNE poll ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Stale data ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | AgeSeconds | Age of the last good board status snapshot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.AgeSeconds | Double | s | Derived | now - Status.LastUpdate | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Bridge freshness contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Stale data ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | Quality | Aggregate snapshot quality: Good, Uncertain, Stale, or Bad. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Quality | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Bridge quality contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Stale data ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | ErrorCount | Number of current status-collection errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.ErrorCount | Integer | N/A | Direct | None | 0 | [13 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=NP04 OPC-UA bridge ErrorCount ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Warning ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Endpoint Status | ErrorsJsonCompatibility | Compatibility JSON list of current status errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.ErrorsJson | String | N/A | Direct | None | [14 empty] | Read-Only | No | On-Change | 5000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Read-only fan-out ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=NP04 OPC-UA bridge ErrorsJson ¦ EXPOSURE=Compatibility only ¦ STATES=All states ¦ ALARM=Status ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls ¦ Typed child nodes are required for production; clients shall not parse JSON for protection logic. | Slow Controls
Authority | ActiveWriter | Current mutation class accepted by the board scheduler: None, DAQ, SlowControls, Expert, or DPSInhibit. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ActiveWriter | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Authority | AuthorityLeaseId | Opaque identifier for the active board-side serialization token, if leases are used. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.AuthorityLeaseId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Authority | AuthorityLeaseExpiresAt | Expiry of the active board-side serialization token, if leases are used. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.AuthorityLeaseExpiresAt | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Authority | DetectorState | Detector/run-control state used by the command gate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.DetectorState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | RunNumber | DAQ run number associated with the active configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.RunNumber | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | RunType | Physics, Noise, Calibration, Commissioning, or Maintenance run classification. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.RunType | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | ConfigurationRevision | Versioned DAQ/PDS configuration revision applied to the board. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ConfigurationRevision | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | ConfigurationHash | Digest of the effective configuration after expansion and defaults. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ConfigurationHash | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | ConfigurationValid | Effective configuration is complete, acknowledged, and valid for the current state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ConfigurationValid | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | DaqConfigureAllowed | All applicable SC/DPS permits, external-activity admission rules, and state guards allow a DAQ-role OPC-UA configuration request. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.DaqConfigureAllowed | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Authority | SlowControlsPermit | Slow-controls operational permit, including applicable site-specific external-activity constraints, presented to the authority gate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.SlowControlsPermit | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | DpsPermit | Detector-protection permit presented to the authority gate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.DpsPermit | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | ExternalConstraintRevision | Version or digest of the approved site-specific external-activity compatibility rules evaluated for this board and detector region. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalConstraintRevision | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Approved site-specific interaction matrix and SC/DPS interface ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | ExternalConstraintStateFresh | All external activity inputs required by the approved compatibility rules are fresh, valid, and applicable to this board and detector region. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalConstraintStateFresh | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Approved site-specific interaction matrix and SC/DPS interface ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | ExternalActivityPermit | SC/DPS admission result indicating whether this board or detector region may be activated under the approved external-activity rules. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalActivityPermit | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Approved site-specific interaction matrix and SC/DPS interface ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | ExternalActivityConflict | Bounded conflict code or list identifying the active/requested external activities that prevent admission; not a substitute for authoritative source states. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalActivityConflict | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Approved site-specific interaction matrix and SC/DPS interface ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | SafeStateActive | Board or upstream supplies are in the agreed safe state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.SafeStateActive | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | MaintenanceLock | Routine DAQ and SC writes are blocked for expert maintenance. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.MaintenanceLock | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | LastOvertakeReason | Reason DAQ authority was removed or superseded by a safe-state/inhibit action. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.LastOvertakeReason | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls / DPS | Slow Controls / DPS
Authority | RecoveryRequired | A fresh DAQ configuration through the OPC-UA gateway is required before data taking can resume. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.RecoveryRequired | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Shared OPC-UA gateway role and field-ownership contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=State guard ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Command Audit | LastCommandId | Opaque identifier of the latest accepted or rejected mutation. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandName | Canonical name of the latest mutating request. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandName | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandClass | DAQ, SC, expert, or protection command class. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandClass | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandRequester | Authenticated user or service identity that requested the mutation. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandRequester | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandAuthority | Authority held when the request was evaluated. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandAuthority | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandRoute | Logical board/mezzanine/AFE/channel route used for the request. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandRoute | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandCorrelationId | Transport correlation identifier for request/reply audit. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandCorrelationId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandRequestTime | UTC source time at request receipt. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandRequestTime | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandCompletionTime | UTC source time at terminal completion. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandCompletionTime | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandState | Rejected, Accepted, Running, Succeeded, Failed, or TimedOut. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandResult | Bounded completion or rejection detail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandResult | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Command Audit | LastCommandReadbackVerified | Required physical/readback postcondition was verified. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandReadbackVerified | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Command broker ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Production command-audit contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Audit ¦ INTERLOCK=No ¦ CONTROL_OWNER=Board service (shared audit) | Board service (shared audit)
Identity | BoardId | Stable DAPHNE board identifier used in the OPC-UA instance path. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.BoardId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB / deployment inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | HardwareDatabaseId | Canonical HWDB component identifier. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.HardwareDatabaseId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | AssetTag | Scanned DAPHNE carrier asset label. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.AssetTag | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB + physical label ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | Detector | Installed detector/module identity. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.Detector | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | DetectorVariant | HD or VD channel/mezzanine profile. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.DetectorVariant | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | GeographicLocation | Controlled geographic name for rack/crate/slot placement. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.GeographicLocation | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | SomProduct | K26 SOM product decoded from EEPROM. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SomProduct | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=K26 EEPROM 0x50 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | SomSerial | K26 SOM serial decoded from EEPROM. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SomSerial | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=K26 EEPROM 0x50 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | SomUuid | K26 SOM UUID decoded from EEPROM. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SomUuid | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=K26 EEPROM 0x50 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | CarrierRevision | DAPHNE carrier revision; no carrier FRU is populated. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.CarrierRevision | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB + asset record ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | FactoryMac | Checksum-valid K26 OEM MAC ID 0. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.FactoryMac | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=K26 EEPROM 0x50 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | ApprovedMac | Network-authority-approved production MAC. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.ApprovedMac | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB / network authority ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | MacIdentityMatch | EEPROM, HWDB, U-Boot, FDT, Linux and network-admission MACs agree. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.MacIdentityMatch | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Derived identity check ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | Hostname | Running Linux hostname. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.Hostname | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Linux + board configuration ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | ExpectedHostname | Hostname assigned by the deployment inventory. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.ExpectedHostname | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=HWDB / deployment inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | HostnameMatch | Runtime hostname equals the assigned value. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.HostnameMatch | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Derived identity check ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Identity | TimingEndpointAddress | Configured DTS endpoint address. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.TimingEndpointAddress | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=endpointControl.ADDRESS / board inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Identity | CrateId | Gateware crate identity field. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.CrateId | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FpgaRegDict.idCrate ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Identity | SlotId | Gateware slot identity field. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SlotId | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FpgaRegDict.idSlot ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Identity | LinkId | Gateware link identity field. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.LinkId | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FpgaRegDict.idLink ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Identity | DetectorId | Gateware detector identity field. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.DetectorId | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FpgaRegDict.idDetector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Identity | RegisterMapVersion | Gateware register-map/interface version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.RegisterMapVersion | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; expected values owned by HWDB/network authority ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FpgaRegDict.idVersion ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=HWDB / network authority; SC monitors | HWDB / network authority; SC monitors
Management Network | Name | Linux network interface name. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Name | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | MacAddress | Active interface MAC address. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.MacAddress | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | PermanentMacAddress | Driver-reported permanent MAC when available. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.PermanentMacAddress | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | Ipv4Address | Active IPv4 address. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Ipv4Address | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | PrefixLength | IPv4 prefix length. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.PrefixLength | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | DefaultGateway | Selected default IPv4 gateway. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.DefaultGateway | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | DnsServers | Configured DNS server list. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.DnsServers | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | VlanId | Assigned VLAN identifier where applicable. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.VlanId | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | OperationalState | Kernel/network-manager operational state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.OperationalState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | Carrier | Physical carrier is present. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Carrier | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | SpeedMbps | Negotiated link speed. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.SpeedMbps | Integer | Mbit/s | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | Duplex | Negotiated duplex mode. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Duplex | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | MtuBytes | Configured interface MTU. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.MtuBytes | Integer | byte | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | RxBytes | Cumulative received bytes. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxBytes | Long | byte | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | RxPackets | Cumulative received packets. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxPackets | Long | packet | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | RxErrors | Cumulative receive errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxErrors | Long | packet | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | RxDropped | Cumulative receive drops. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxDropped | Long | packet | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | TxBytes | Cumulative transmitted bytes. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxBytes | Long | byte | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | TxPackets | Cumulative transmitted packets. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxPackets | Long | packet | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | TxErrors | Cumulative transmit errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxErrors | Long | packet | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | TxDropped | Cumulative transmit drops. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxDropped | Long | packet | Direct | None | [14 empty] | Read-Only | Yes | 10000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | LinkChangeCount | Observed carrier/operational-state transitions. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.LinkChangeCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Management Network | ConfigurationMatch | MAC/IP/prefix/gateway/VLAN match the deployment inventory. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.ConfigurationMatch | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; network assignment owned externally ¦ INSTANCE=per interface; baseline instance eth0 ¦ SOURCE=Linux rtnetlink/sysfs plus HWDB/network inventory ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Network health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | CurrentUtc | Board UTC clock. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.CurrentUtc | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Linux clock ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | UnixTimeNs | Board UNIX time in nanoseconds. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.UnixTimeNs | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Linux clock ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | UptimeSeconds | Host uptime. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.UptimeSeconds | Long | s | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Linux /proc/uptime ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | BootId | Linux boot identifier. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.BootId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Linux /proc/sys/kernel/random/boot_id ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | BootTime | Derived host boot time. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.BootTime | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=CurrentUtc - UptimeSeconds ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | KernelRelease | Running kernel release. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.KernelRelease | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=uname ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | OperatingSystemVersion | PetaLinux/OS release. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.OperatingSystemVersion | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=/etc/os-release ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | RootFilesystemBuildId | Installed rootfs/release build identity. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.RootFilesystemBuildId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=release manifest ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | ActiveBootSlot | Active A/B eMMC runtime slot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.ActiveBootSlot | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=boot-state service/U-Boot env ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | BootMarkedGood | Current runtime slot has passed the boot-success contract. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.BootMarkedGood | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=daphne-boot-ok.service ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | LastResetReason | Best available reboot/reset reason. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.LastResetReason | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=kernel/PMU/boot evidence ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | WatchdogHealthy | Host watchdog is configured and being serviced. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.WatchdogHealthy | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=watchdog/sysfs ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | NtpSynchronized | System clock is synchronized to the approved time source. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.NtpSynchronized | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timesync service ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | NtpOffsetMilliseconds | Reported clock offset from the approved source. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.NtpOffsetMilliseconds | Double | ms | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timesync service ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | TimeSource | Active time synchronization source. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.TimeSource | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timesync service ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | CpuLoad1Minute | One-minute CPU load average. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.CpuLoad1Minute | Double | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=/proc/loadavg ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | MemoryAvailableBytes | Available system memory. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.MemoryAvailableBytes | Long | byte | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=/proc/meminfo ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | RootFilesystemFreeBytes | Free space on the runtime root filesystem. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.RootFilesystemFreeBytes | Long | byte | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 60000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=statvfs ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Host | RootFilesystemReadOnly | Runtime root filesystem is unexpectedly read-only. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.RootFilesystemReadOnly | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=mount state ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Host health ¦ INTERLOCK=No ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Services | Active | Service is in the active state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.Active | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; service recovery owned by PDS operations ¦ INSTANCE=per required service; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.ServiceStatus / systemd ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Service fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Services | State | Service active/substate detail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.State | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; service recovery owned by PDS operations ¦ INSTANCE=per required service; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.ServiceStatus / systemd ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Service fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Services | MainPid | Main process identifier, if applicable. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.MainPid | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; service recovery owned by PDS operations ¦ INSTANCE=per required service; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.ServiceStatus / systemd ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Service fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Services | RestartCount | Cumulative automatic restart count for this boot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.RestartCount | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; service recovery owned by PDS operations ¦ INSTANCE=per required service; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.ServiceStatus / systemd ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Service fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Services | StartTime | Time the current service instance started. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.StartTime | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; service recovery owned by PDS operations ¦ INSTANCE=per required service; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.ServiceStatus / systemd ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Service fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Services | LastExitStatus | Most recent terminal exit status/result. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.LastExitStatus | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; service recovery owned by PDS operations ¦ INSTANCE=per required service; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.ServiceStatus / systemd ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Service fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
OPC-UA Bridge | Heartbeat | Monotonic OPC-UA bridge heartbeat. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.Heartbeat | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | Version | Running bridge software version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.Version | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | ConfigurationPath | Controlled configuration source path or identifier. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ConfigurationPath | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | ConfigurationHash | Digest of the effective bridge configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ConfigurationHash | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | NamespaceVersion | Published OPC-UA information-model version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.NamespaceVersion | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | BoardCount | Number of configured DAPHNE board instances. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.BoardCount | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | WorkerCount | Bounded background DAPHNE poll worker count. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.WorkerCount | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | CertificateNotAfter | OPC-UA server certificate expiry. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.CertificateNotAfter | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | RejectedSessionCount | Cumulative rejected OPC-UA sessions. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.RejectedSessionCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | PollTimeoutCount | Cumulative DAPHNE polling timeouts. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.PollTimeoutCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | PollErrorCount | Cumulative DAPHNE polling errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.PollErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | BackendProtocolVersion | DAPHNE Protobuf/ZMQ backend schema and compatibility-adapter version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.BackendProtocolVersion | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | BackendCompatibilityState | Compatibility of the gateway adapter with every configured board: Compatible, Degraded, or Incompatible. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.BackendCompatibilityState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | ActiveDaqSessionCount | Authenticated active OPC-UA sessions holding the DAQ client role. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ActiveDaqSessionCount | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | ActiveSlowControlsSessionCount | Authenticated active OPC-UA sessions holding the Slow Controls client role. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ActiveSlowControlsSessionCount | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | UnauthorizedMethodCount | Cumulative method calls rejected by OPC-UA role or field-level authorization. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.UnauthorizedMethodCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | CommandQueueDepth | Number of accepted board mutations waiting in the gateway serializer. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.CommandQueueDepth | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
OPC-UA Bridge | TranslationErrorCount | Cumulative failures translating typed OPC-UA operations to or from the DAPHNE backend protocol. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.TranslationErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Neutral PDS gateway deployment shared by DAQ and SC ¦ INSTANCE=one per bridge shard ¦ SOURCE=NP04 bridge for current fields; production shared-gateway contract otherwise ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Bridge health ¦ INTERLOCK=No ¦ CONTROL_OWNER=PDS OPC-UA gateway service | PDS OPC-UA gateway service
Firmware | Loaded | Expected FPGA image is loaded and FPGA manager reports operating. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.Loaded | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.FirmwareStatus.loaded ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | FpgaManagerState | Raw FPGA manager state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FpgaManagerState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FirmwareStatus.fpga_manager_state ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | OverlayName | Active FPGA application/overlay name. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.OverlayName | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FirmwareStatus.overlay_name ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | BuildId | Gateware build identifier extracted from release metadata. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.BuildId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FirmwareStatus.build_id ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | GitCommit | Gateware Git revision read from the configuration register. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GitCommit | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FpgaRegDict.GIT ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | ExpectedBuildId | Build identity required by the active release/configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ExpectedBuildId | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=release manifest ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | BuildMatchesExpected | Running FPGA build matches the release/configuration contract. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.BuildMatchesExpected | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=derived comparison ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | RegisterMapVersionMatches | Gateware register-map version matches the board-service schema. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.RegisterMapVersionMatches | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=idVersion + server manifest ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | PlI2cDevicePresent | PL I2C platform device is present. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.PlI2cDevicePresent | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FirmwareStatus.pl_devices_present / preflight ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | PlSpiDevicePresent | PL SPI platform device is present. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.PlSpiDevicePresent | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=FirmwareStatus.pl_devices_present / preflight ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | ActiveXmutilApplication | Active xmutil accelerator application. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ActiveXmutilApplication | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=xmutil listapps ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | FirmwareUpdateState | Idle, Preparing, Programming, Verifying, Succeeded, or Failed. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FirmwareUpdateState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=firmware update contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | FirmwarePathAvailable | Normal configuration/monitor path remains available. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FirmwarePathAvailable | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=firmware update contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | FirmwareArtifactSha256 | Digest of the loaded/installed FPGA artifact. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FirmwareArtifactSha256 | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=release/deployment manifest ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | ServerVersion | DAPHNE board-service software version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ServerVersion | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=board service build metadata ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | ProtobufSchemaVersion | ControlEnvelope/Protobuf schema compatibility version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ProtobufSchemaVersion | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=board service build metadata ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | GatewareHeartbeat | Monotonic PL heartbeat or watchdog counter. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GatewareHeartbeat | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required gateware health block ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | GatewareFatalFaultLatched | Latched fatal gateware error state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GatewareFatalFaultLatched | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required gateware health block ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | GatewareResetReason | Last gateware reset cause. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GatewareResetReason | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required gateware health block ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Firmware | ConfigurationReady | Analog/front-end configuration path has reached a stable ready state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ConfigurationReady | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; updates are expert maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=analog-control contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Firmware readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | Available | RPU/remoteproc control backend is available. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Available | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | Running | RPU firmware is running. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Running | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Yes ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | Firmware | RPU firmware image name/version. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Firmware | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | Heartbeat | Monotonic RPU safety/control heartbeat. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Heartbeat | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Yes ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | LastFault | Most recent RPU fault detail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.LastFault | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | AbiVersion | RPU wire ABI version accepted by the Linux board service. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.AbiVersion | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | InterlockReady | RPU-side voltage/current/temperature/communication preconditions allow AFE writes. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.InterlockReady | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Yes ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | FaultCode | Structured current RPU fault bitmap/code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.FaultCode | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Yes ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | CommandQueueDepth | Bounded RPU command queue occupancy. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.CommandQueueDepth | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls | Slow Controls
RPU | LastReplySequence | Sequence number of the most recent completed RPU command. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.LastReplySequence | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; RPU owns safety-critical AFE sequencing ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=SlowControlStatusResponse.RpuStatus / RPU wire status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Control-path fault ¦ INTERLOCK=Context ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | Present | Linux I2C adapter device node exists. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Present | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | Path | Linux I2C adapter path. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Path | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | Functional | A bounded non-disruptive health transaction succeeds. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Functional | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | Owner | Process/RPU component currently serializing access. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Owner | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | Busy | A serialized I2C transaction is in progress. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Busy | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | LastSuccess | Time of the last successful health transaction. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.LastSuccess | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | ErrorCount | Cumulative transaction errors for this boot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.ErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | TimeoutCount | Cumulative transaction timeouts for this boot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.TimeoutCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Bus | RecoveryCount | Cumulative bus recoveries for this boot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.RecoveryCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized hardware owner ¦ INSTANCE=per I2C bus; expected buses 1 and 2 ¦ SOURCE=SlowControlStatusResponse.I2CBusStatus plus required metrics ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Device | Name | Stable device name. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Name | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per expected and discovered I2C device; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.I2CDeviceStatus / device driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C device fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Device | Present | Expected I2C address responds/is registered. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Present | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per expected and discovered I2C device; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.I2CDeviceStatus / device driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C device fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Device | Driver | Bound Linux driver or board-service driver. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Driver | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=per expected and discovered I2C device; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.I2CDeviceStatus / device driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C device fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Device | Status | Bounded device-specific health detail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per expected and discovered I2C device; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.I2CDeviceStatus / device driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C device fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Device | LastSuccess | Time of the last successful device transaction. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.LastSuccess | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=per expected and discovered I2C device; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.I2CDeviceStatus / device driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C device fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
I2C Device | ErrorCount | Cumulative device transaction errors for this boot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.ErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 10000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=per expected and discovered I2C device; see Instance Sets ¦ SOURCE=SlowControlStatusResponse.I2CDeviceStatus / device driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=I2C device fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SPI | Present | PL SPI device and /dev/spidev3.0 are present. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.Present | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; one serialized owner ¦ INSTANCE=per SPI device; baseline current monitor ¦ SOURCE=preflight /dev/spidev3.0 and current-monitor driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=SPI health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SPI | Functional | Bounded current-monitor SPI identity/read transaction succeeds. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.Functional | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized owner ¦ INSTANCE=per SPI device; baseline current monitor ¦ SOURCE=preflight /dev/spidev3.0 and current-monitor driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=SPI health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SPI | Owner | Component currently serializing SPI access. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.Owner | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized owner ¦ INSTANCE=per SPI device; baseline current monitor ¦ SOURCE=preflight /dev/spidev3.0 and current-monitor driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=SPI health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SPI | ErrorCount | Cumulative SPI transaction errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.ErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized owner ¦ INSTANCE=per SPI device; baseline current monitor ¦ SOURCE=preflight /dev/spidev3.0 and current-monitor driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=SPI health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SPI | LastSuccess | Time of the last successful SPI transaction. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.LastSuccess | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; one serialized owner ¦ INSTANCE=per SPI device; baseline current monitor ¦ SOURCE=preflight /dev/spidev3.0 and current-monitor driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=SPI health ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 3V3PDSVoltage | Measured 3V3PDS board rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3PDS.Voltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 3V3PDSStatus | Quality/fault state for the 3V3PDS board rail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3PDS.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required typed rail status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 1V8PDSVoltage | Measured 1V8PDS board rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8PDS.Voltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 1V8PDSStatus | Quality/fault state for the 1V8PDS board rail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8PDS.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required typed rail status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 3V3AVoltage | Measured 3V3A board rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3A.Voltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 3V3AStatus | Quality/fault state for the 3V3A board rail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3A.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required typed rail status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 1V8AVoltage | Measured 1V8A board rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8A.Voltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | 1V8AStatus | Quality/fault state for the 1V8A board rail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8A.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required typed rail status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | Minus5VAVoltage | Measured Minus5VA board rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.Minus5VA.Voltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Board Rails | Minus5VAStatus | Quality/fault state for the Minus5VA board rail. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.Minus5VA.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required typed rail status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | Voltage | Measured regulator output voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Voltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | Current | Measured regulator output current. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Current | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | Power | Measured/derived regulator output power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Power | Double | W | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | Temperature | Regulator temperature readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Temperature | Double | °C | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | StatusByte | Raw PMBus STATUS_BYTE. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusByte | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | StatusWord | Raw PMBus STATUS_WORD. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusWord | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | StatusVout | PMBus output-voltage fault bitmap. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusVout | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | StatusIout | PMBus output-current fault bitmap. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusIout | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | StatusTemperature | PMBus temperature fault bitmap. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusTemperature | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | StatusCml | PMBus communication/memory/logic fault bitmap. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusCml | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
PMBus Regulators | Health | Decoded regulator health summary. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Health | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 ¦ SOURCE=PJT004A0X43 / Rust PMBus collector ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Rail fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Temperature | Celsius | Typed temperature reading. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Sensors.{Sensor}.Celsius | Double | °C | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only ¦ INSTANCE=per discovered allow-listed sensor ¦ SOURCE=SlowControlStatusResponse.TemperatureReading; Linux thermal/hwmon; PMBus/MCP9808/SFP ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Thermal fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Temperature | Status | Good, Warning, High, Critical, Missing, or Invalid. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Sensors.{Sensor}.Status | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=per discovered allow-listed sensor ¦ SOURCE=SlowControlStatusResponse.TemperatureReading; Linux thermal/hwmon; PMBus/MCP9808/SFP ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Thermal fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Temperature | LastUpdate | Source time of the temperature reading. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Sensors.{Sensor}.LastUpdate | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only ¦ INSTANCE=per discovered allow-listed sensor ¦ SOURCE=SlowControlStatusResponse.TemperatureReading; Linux thermal/hwmon; PMBus/MCP9808/SFP ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Thermal fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | BlockEnabled | HD mezzanine I2C block is selected/enabled. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.BlockEnabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Configured | HD mezzanine monitor/protection configuration is valid. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Configured | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Power5VRequested | Requested 5 V enable state from the paired power command. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power5VRequested | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Power3V3Requested | Requested 3.3 V enable state from the paired power command. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power3V3Requested | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Voltage5V | Measured 5 V rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Voltage5V | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Current5V | Measured 5 V rail current. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Current5V | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Power5V | Measured 5 V rail power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power5V | Double | W | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Alert5V | Latched INA232 5 V alert. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Alert5V | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Voltage3V3 | Measured 3.3 V rail voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Voltage3V3 | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Current3V3 | Measured 3.3 V rail current. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Current3V3 | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Power3V3 | Measured 3.3 V rail power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power3V3 | Double | W | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | Alert3V3 | Latched INA232 3.3 V alert. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Alert3V3 | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | RShunt5V | Configured 5 V shunt resistance. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.RShunt5V | Double | Ohm | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | RShunt3V3 | Configured 3.3 V shunt resistance. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.RShunt3V3 | Double | Ohm | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | MaxCurrentScale5V | 5 V measurement full-scale current configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxCurrentScale5V | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | MaxCurrentScale3V3 | 3.3 V measurement full-scale current configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxCurrentScale3V3 | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | ShutdownCurrent5V | 5 V hardware alert/shutdown-current configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShutdownCurrent5V | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | ShutdownCurrent3V3 | 3.3 V hardware alert/shutdown-current configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShutdownCurrent3V3 | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | MaxPower5V | Derived 5 V maximum measurable power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxPower5V | Double | W | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | MaxPower3V3 | Derived 3.3 V maximum measurable power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxPower3V3 | Double | W | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | CurrentLsb5V | Configured INA232 5 V current LSB. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.CurrentLsb5V | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | CurrentLsb3V3 | Configured INA232 3.3 V current LSB. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.CurrentLsb3V3 | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | ShuntCal5V | INA232 5 V calibration register readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShuntCal5V | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
HD Mezzanine | ShuntCal3V3 | INA232 3.3 V calibration register readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShuntCal3V3 | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=SC owns rail requests; monitor values are read-only ¦ INSTANCE=per AFE block 0..4 when HD mezzanine is installed ¦ SOURCE=HD mezzanine Protobuf API and cached INA232/TCA9536 driver ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Mezzanine power fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | Present | SFP module is present. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.Present | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | LossOfSignal | Receiver loss-of-signal input is asserted. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.LossOfSignal | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | TxFault | SFP transmitter fault input is asserted. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxFault | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | TxDisabled | SFP transmitter is disabled. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxDisabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | RateSelect | SFP rate-select state when wired. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.RateSelect | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | VendorName | SFP vendor name from EEPROM. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.VendorName | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | VendorOui | SFP vendor OUI. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.VendorOui | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | PartNumber | SFP vendor part number. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.PartNumber | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | Revision | SFP vendor revision. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.Revision | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | SerialNumber | SFP serial number. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.SerialNumber | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | DateCode | SFP manufacturing date code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.DateCode | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | NominalBitRateMbps | SFP nominal signaling rate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.NominalBitRateMbps | Integer | Mbit/s | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | WavelengthNm | SFP nominal transmit wavelength. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.WavelengthNm | Double | nm | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFF-8472 plus EDMS 2580777 endpoint SFP specification ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls ¦ For {Sfp}=timing, the authoritative endpoint specification is TX 1550 nm / RX 1490 nm. Other SFP roles use their approved role-specific inventory values. | Slow Controls
SFP | TemperatureCelsius | SFF-8472 module temperature. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TemperatureCelsius | Double | °C | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | SupplyVoltage | SFF-8472 module supply voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.SupplyVoltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | TxBiasCurrent | SFF-8472 laser bias current. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxBiasCurrent | Double | mA | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | TxOpticalPower | SFF-8472 transmit optical power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxOpticalPower | Double | mW | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | RxOpticalPower | SFF-8472 receive optical power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.RxOpticalPower | Double | mW | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | AlarmFlags | Decoded SFF-8472 alarm flags. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.AlarmFlags | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | WarningFlags | Decoded SFF-8472 warning flags. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.WarningFlags | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | DomSupported | Digital optical monitoring is supported. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.DomSupported | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
SFP | EepromReadable | SFP identity/diagnostic EEPROM is readable. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.EepromReadable | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; TX disable controlled by endpoint/link firmware ¦ INSTANCE=per timing, CCM, and provisioned data SFP ¦ SOURCE=SFP GPIO/I2C expander 0x72 and SFF-8472 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Optical link fault ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | Present | Fan is installed/detected. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.Present | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | PwmCommand | Applied fan PWM/control code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.PwmCommand | Integer | count | Direct | None | 0 | 255 | [12 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | TachometerRaw | Raw FPGA fan-speed counter. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.TachometerRaw | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | SpeedRpm | Fan speed after the validated tachometer conversion. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.SpeedRpm | Double | rpm | Direct | None | [14 empty] | Read-Only | Yes | 1000 | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | Running | Fan tachometer indicates rotation. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.Running | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | Stalled | Commanded fan is below the approved minimum speed. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.Stalled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | StallCount | Cumulative detected stalls for this boot. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.StallCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | ControlMode | Automatic thermal control or maintenance override. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.ControlMode | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Fans | LastUpdate | Source time of the latest fan sample. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.LastUpdate | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Local automatic control; maintenance override only ¦ INSTANCE=per fan 0..1 ¦ SOURCE=FpgaRegDict fanControl/fanReadSpeed_0..1 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Fan fault ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
Timing | ClockSourceControlled | Endpoint clock source is under the configured timing-control path. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.ClockSourceControlled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=endpointClockControl.CLOCK_SOURCE ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | Mmcm0Locked | Timing MMCM0 lock state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.Mmcm0Locked | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ClockStatus.mmcm0_locked ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | Mmcm1Locked | Timing MMCM1 lock state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.Mmcm1Locked | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ClockStatus.mmcm1_locked ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | EndpointFsmState | Decoded timing endpoint FSM state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointFsmState | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=endpointStatus.FSM_STATUS ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | EndpointFsmRaw | Raw timing endpoint FSM code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointFsmRaw | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=endpointStatus.FSM_STATUS ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | TimestampValid | Endpoint timestamp-valid bit. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TimestampValid | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=endpointStatus.TIMESTAMP_OK ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | TimingUsable | Derived readiness: required locks, accepted FSM state, and timestamp valid. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TimingUsable | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=preflight readiness contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | RawEndpointStatus | Raw endpoint status register. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.RawEndpointStatus | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ClockStatus.raw_endpoint_status ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | EndpointAddress | Timing endpoint address readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointAddress | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=endpointControl.ADDRESS ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | TxEnableRequested | DTS/endpoint logic has requested transmitter enable. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TxEnableRequested | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timing endpoint contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | TxDisabled | Timing SFP TX_DISABLE output state. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TxDisabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timing endpoint output ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | LossOfSignal | Timing SFP LOS observation. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LossOfSignal | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=sfp_tmg_los ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | LockLossCount | Cumulative endpoint lock-loss events. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LockLossCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required timing health block ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | AutoRelockCount | Cumulative automatic relock attempts. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.AutoRelockCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required timing health block ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | LastLockLossTime | Time of the most recent timing lock loss. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LastLockLossTime | DateTime | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=required timing health block ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | Timestamp | Current assembled DTS timestamp sample. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.Timestamp | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timestamp0..3 register words ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | LastDtsCommandId | Most recent synchronous DTS command identifier observed. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LastDtsCommandId | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=timing/header integration contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | GroupMask | Applied timing group-address mask. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.GroupMask | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=versioned timing configuration ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | DelayTrim | Compatibility readback of the applied endpoint delay compensation. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.DelayTrim | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=versioned timing configuration ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | CommandDecoderMap | Applied mapping of DTS command codes to decoder outputs 0..7. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.CommandDecoderMap | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=EDMS 2395364 v2 control-bus registers 0x60..0x67 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | ClockPhaseUpdateComplete | Endpoint reports completion of the latest clock-phase update. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.ClockPhaseUpdateComplete | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=EDMS 2395364 v2 status register 0x71 bit 4 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | CoarseDelay | Applied endpoint coarse deskew delay. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.CoarseDelay | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=EDMS 2395364 v2 deskew registers 0x72..0x73 bits 3..0 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | FinePhase | Applied endpoint fine clock-phase setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.FinePhase | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=EDMS 2395364 v2 deskew registers 0x72..0x73 bits 15..4 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | ProtocolErrorCount | Cumulative soft DTS protocol-error count. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.ProtocolErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=EDMS 2395364 v2 err_cnt register 0x76 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
Timing | EndpointVersion | Timing endpoint firmware-interface version readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointVersion | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=EDMS 2395364 v2 version registers 0x78..0x7b ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Timing readiness ¦ INTERLOCK=Readiness ¦ CONTROL_OWNER=DAQ / DTS | DAQ / DTS
AFE Global | PowerState | AFE global power-state readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.PowerState | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=SC owns power/bias enable and safe-state action ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=afeGlobalControl.POWERSTATE ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
AFE Global | ResetAsserted | AFE global reset readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.ResetAsserted | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ owns data-taking configuration/reset; SC monitors ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=afeGlobalControl.RESET ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=DAQ | DAQ
AFE Global | BusyAfe0 | AFE0 configuration busy. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BusyAfe0 | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ owns data-taking configuration/reset; SC monitors ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=afeGlobalControl.BUSY_AFE0 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=DAQ | DAQ
AFE Global | BusyAfe12 | AFE1/2 configuration busy. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BusyAfe12 | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ owns data-taking configuration/reset; SC monitors ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=afeGlobalControl.BUSY_AFE12 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=DAQ | DAQ
AFE Global | BusyAfe34 | AFE3/4 configuration busy. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BusyAfe34 | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ owns data-taking configuration/reset; SC monitors ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=afeGlobalControl.BUSY_AFE34 ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=DAQ | DAQ
AFE Global | BiasEnable | Board bias-enable register readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BiasEnable | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=LOCAL READABLE ¦ AUTHORITY=SC owns power/bias enable and safe-state action ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=biasEnable.ENABLE ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=Slow Controls | Slow Controls
AFE Global | BiasControlCode | Board VBIAS-control DAC/code readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BiasControlCode | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ owns data-taking configuration/reset; SC monitors ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=READ_VBIAS_CONTROL ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE readiness ¦ INTERLOCK=Candidate ¦ CONTROL_OWNER=DAQ | DAQ
AFE | Attenuation | Applied AFE attenuation code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.Attenuation | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=READ/WRITE AFE attenuation ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | VGain | Applied AFE VGAIN code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.VGain | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=READ_AFE_VGAIN ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasSetCode | Applied AFE bias DAC set code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasSetCode | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=READ_AFE_BIAS_SET ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasVoltage | Measured AFE bias voltage. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltage | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT OPC-UA POC ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=GeneralInfo.v_bias_0..4 / bias monitor ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasVoltageTarget | Requested physical closed-loop bias target. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltageTarget | Double | V | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=cmd_writeAFEBiasControlledSet.target_bias_mv ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasVoltageError | Final measured-minus-target bias error. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltageError | Double | mV | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=cmd_writeAFEBiasControlledSet_response.error_mv ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasControlStatus | Closed-loop bias outcome enum. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasControlStatus | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AfeBiasControlStatus ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasControlSettled | Closed-loop bias reached its settling criterion. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasControlSettled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=controlled-set response.settled ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasControlIterations | Closed-loop bias iterations used. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasControlIterations | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=controlled-set response.iterations ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasVoltageSlope | Final measured bias slope. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltageSlope | Double | mV/s | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=controlled-set response.final_slope_mv_per_s ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | BiasStableSamples | Consecutive stable samples at completion. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasStableSamples | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=controlled-set response.stable_samples ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | AdcResolution | Applied ADC resolution setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AdcResolution | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.adc.resolution ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | AdcOutputFormat | Applied ADC output-format setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AdcOutputFormat | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.adc.output_format ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | AdcMsbFirst | Applied ADC serialization bit-order setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AdcMsbFirst | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.adc.sb_first ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | PgaLpfCutFrequency | Applied PGA low-pass filter code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.PgaLpfCutFrequency | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.pga.lpf_cut_frequency ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | PgaIntegratorDisabled | Applied PGA integrator-disable setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.PgaIntegratorDisabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.pga.integrator_disable ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | PgaGain | Applied PGA gain setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.PgaGain | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.pga.gain ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | LnaClamp | Applied LNA clamp code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.LnaClamp | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.lna.clamp ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | LnaGain | Applied LNA gain code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.LnaGain | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.lna.gain ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | LnaIntegratorDisabled | Applied LNA integrator-disable setting. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.LnaIntegratorDisabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=AFEConfig.lna.integrator_disable ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | AlignmentDelay | Selected IDELAY tap. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AlignmentDelay | Integer | tap | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=cmd_alignAFEs_response.delay ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | AlignmentBitslip | Selected bitslip value. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AlignmentBitslip | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=cmd_alignAFEs_response.bitslip ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
AFE | Aligned | AFE alignment has a valid matched result. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.Aligned | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits ¦ INSTANCE=per AFE block 0..4 ¦ SOURCE=frontend alignment status contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=AFE configuration/readback ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ | DAQ
Channels | Trim | Applied per-channel trim code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Trim | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=READ_TRIM_* ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | Offset | Applied per-channel offset code. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Offset | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=READ_OFFSET_* ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | Gain | Applied per-channel gain selection/code from the active configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Gain | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=ChannelConfig.gain + readback contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | CurrentMonitorRaw | Raw per-channel current-monitor ADC result. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.CurrentMonitorRaw | Integer | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=READ_CURRENT_MONITOR ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | Current | Calibrated per-channel current monitor value. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Current | Double | A | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=current monitor + calibration ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | TriggerEnabled | Effective per-channel trigger-enable mask bit. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerEnabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=triggerEnableLow/High ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | FullStreamEnabled | Channel is selected for full-stream data. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.FullStreamEnabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=ConfigureRequest.full_stream_channels ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | TriggerThreshold | Applied 10-bit self-trigger threshold. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerThreshold | Integer | count | Direct | None | 0 | 1023 | [12 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=TriggerChannelSnapshot.threshold ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | TriggerRecordCount | Cumulative self-trigger record count. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerRecordCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 30000 | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=TriggerChannelSnapshot.record_count ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | TriggerBusyCount | Cumulative self-trigger busy count. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerBusyCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 30000 | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=TriggerChannelSnapshot.busy_count ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | TriggerFullCount | Cumulative self-trigger full count. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerFullCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 30000 | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=TriggerChannelSnapshot.full_count ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | TriggerRate | Derived trigger-record rate over the polling interval. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerRate | Double | Hz | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=delta(record_count)/delta(time) ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | BusyFraction | Derived busy-count fraction over the polling interval. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.BusyFraction | Double | % | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=delta(busy)/defined opportunity denominator ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | PacketCount | Cumulative per-channel packet count. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.PacketCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 30000 | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=ChannelInfo.total_packets ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Channels | PacketRate | Derived per-channel packet rate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.PacketRate | Double | Hz | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set ¦ INSTANCE=per channel; HD 0..39, VD 0..31 ¦ SOURCE=ChannelInfo.new_packets / interval ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Run context ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | SelfTriggerThreshold | Applied aggregate/self-trigger threshold configuration word. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerThreshold | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.self_trigger_threshold ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | SelfTriggerXcorr | Applied cross-correlation trigger configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerXcorr | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.self_trigger_xcorr ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | TriggerPrimitiveConfig | Applied trigger-primitive configuration word. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.TriggerPrimitiveConfig | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.tp_conf ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | CompensatorMask | Applied AFE/channel compensator mask. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.CompensatorMask | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.compensator ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | InverterMask | Applied AFE/channel inverter mask. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.InverterMask | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.inverters ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | FullStreamChannelList | Applied full-stream channel list. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.FullStreamChannelList | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.full_stream_channels ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | Slot | Applied DAQ slot identity. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.Slot | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.slot ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | RequestTimeoutMs | Applied bounded command timeout. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.RequestTimeoutMs | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.timeout_ms ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | BiasControl | Applied board bias-control configuration word. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.BiasControl | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=ConfigureRequest.biasctrl ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | SelfTriggerConfigRaw | Raw self-trigger configuration register. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerConfigRaw | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=stuffregs.OFF_ST_CONFIG ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | SelfTriggerDelayRaw | Raw self-trigger delay register. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerDelayRaw | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=stuffregs.OFF_ST_DELAY ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Run Configuration | SelfTriggerFilterOutputSelect | Raw self-trigger filter output selection. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerFilterOutputSelect | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 30000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=stuffregs.OFF_ST_FILTER_OUTPUT_SEL ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Configuration mismatch ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Spy Trigger | SourceSelector | Selected spy-buffer trigger source: 0 Software, 1 External, 2 TimingAdHoc, or 3 LegacyOrAll. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Spy.Trigger.SourceSelector | Integer | N/A | Direct | decode FRONT_END_S_AXI+0x34 bits [1:0] | 0 | 3 | 3 | 0 | 3 | N/A (categorical) | N/A (categorical) | N/A (categorical) | N/A (categorical) | N/A (categorical) | N/A (categorical) | 0 | 5 | No protection interlock. Reject values outside 0..3; if the readback differs from the active Run Control profile for 5 s, alarm and mark the spy-buffer configuration mismatched. | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ OPC-UA configuration; SC monitoring ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=daphne-firmware working tree: spy_trigger_ctrl_reg[1:0] at 0x88000034 ¦ EXPOSURE=Baseline DCS readback; write only through guarded typed operation ¦ STATES=All states ¦ ALARM=Spy-buffer configuration ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ ¦ ENUM=0 Software; 1 External; 2 TimingAdHoc; 3 LegacyOrAll. RESET/NOMINAL=3 for backward compatibility; the operational target is supplied directly by the active DAQ profile. This is distinct from SelfTriggerFilterOutputSelect at 0x94000034. | DAQ
Spy Trigger | Inhibit | Spy-buffer trigger inhibit; true blocks software, external, and timing/ad-hoc triggers before capture. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Spy.Trigger.Inhibit | Boolean | N/A | Direct | decode FRONT_END_S_AXI+0x34 bit [2] | 0 | 1 | FALSE | FALSE | FALSE | N/A (Boolean) | N/A (Boolean) | N/A (Boolean) | TRUE | N/A (Boolean) | N/A (not a DPS trip) | 0 | 5 | No DPS hard-interlock action. If true outside a guarded source change or approved inhibited state for 5 s, alarm and mark spy-buffer capture not ready. | Read-Only | Yes | On-Change | 5000 | STATUS=LOCAL READABLE ¦ AUTHORITY=DAQ OPC-UA configuration; SC monitoring ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=daphne-firmware working tree: spy_trigger_ctrl_reg[2] at 0x88000034 ¦ EXPOSURE=Baseline DCS readback; write only through guarded typed operation ¦ STATES=All states ¦ ALARM=Spy capture inhibited ¦ INTERLOCK=Context ¦ CONTROL_OWNER=DAQ ¦ RESET/NOMINAL=False. True is allowed transiently during an atomic source change and may be held only by an approved DAQ/expert configuration. This inhibits spy-buffer capture triggers only and is not a DPS hard inhibit. | DAQ
Readout Stream | TotalPackets | Cumulative packets offered/sent by the DAPHNE stream. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.TotalPackets | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | NewPackets | Packets observed during the latest monitoring interval. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.NewPackets | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | TotalDroppedPackets | Cumulative dropped packets. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.TotalDroppedPackets | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | NewDroppedPackets | Dropped packets in the latest monitoring interval. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.NewDroppedPackets | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | PacketRate | Derived packet rate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.PacketRate | Double | Hz | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | DropRate | Derived dropped-packet rate. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.DropRate | Double | Hz | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | DropFraction | Derived fraction of offered packets dropped. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.DropFraction | Double | % | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | ThrottlingActive | Readout throttling is currently active. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.ThrottlingActive | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | ThrottleEventCount | Cumulative throttling events. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.ThrottleEventCount | Long | count | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Readout Stream | DataPathStale | Readout datapath heartbeat/counters are not advancing as expected. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.DataPathStale | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | 5000 | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Read-only; DAQ owns data-path recovery ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=InfoResponse.StreamInfo plus required gateware/Hermes health ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout readiness ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | Provisioned | 10 GbE lane is provisioned for this board. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.Provisioned | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | Enabled | Lane transmitter is enabled. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.Enabled | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | LinkReady | Lane clock/reset/PCS path is ready. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.LinkReady | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | LocalFault | 10 GbE local fault indication. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.LocalFault | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | RemoteFault | 10 GbE remote fault indication. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.RemoteFault | Boolean | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | TxPacketCount | Cumulative packets on this lane. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.TxPacketCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | TxByteCount | Cumulative bytes on this lane. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.TxByteCount | Long | byte | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | ErrorCount | Cumulative lane/PCS errors. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.ErrorCount | Long | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | DestinationIp | Configured DAQ destination IP. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.DestinationIp | String | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | DestinationPort | Configured DAQ destination UDP port. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.DestinationPort | Integer | N/A | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
10GbE Readout | MtuBytes | Configured jumbo-frame MTU. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.MtuBytes | Integer | byte | Direct | None | [14 empty] | Read-Only | Yes | On-Change | 5000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback ¦ INSTANCE=per provisioned 10 GbE data lane; hardware supports up to four ¦ SOURCE=Hermes/10GbE gateware health contract ¦ EXPOSURE=Baseline DCS ¦ STATES=All states ¦ ALARM=Readout link fault ¦ INTERLOCK=No ¦ CONTROL_OWNER=DAQ | DAQ
Operation | RequestSafeState | Method: Request the agreed DAPHNE safe state without supplying a competing run configuration. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.RequestSafeState | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Slow Controls/DPS ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=Baseline controlled ¦ STATES=Normal, Running, Fault ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls / DPS ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls / DPS
Operation | AcknowledgeRecovery | Method: Acknowledge SC recovery conditions; DAQ must still perform a fresh configure. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.AcknowledgeRecovery | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Slow Controls ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=Baseline controlled ¦ STATES=Recovery ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | ConfigureRun | Method: Apply the complete versioned DAPHNE run configuration from the authenticated DAQ OPC-UA role. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureRun | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only; not an Ignition operator button ¦ STATES=Configure/Calibration ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | ConfigureClocks | Method: Apply endpoint address/clock-source configuration and bounded resets. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureClocks | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ/DTS; PDS expert in maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ/DTS OPC-UA roles only ¦ STATES=Configure/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ / DTS ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ / DTS
Operation | ConfigureFrontend | Method: Apply DAQ-owned AFE and per-channel configuration atomically/staged. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureFrontend | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only; not a casual operator write ¦ STATES=Configure/Calibration ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | SetControlledBias | Method: Set the DAQ-owned numerical AFE bias target using the bounded closed-loop controller; SC separately owns bias enable and permits. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetControlledBias | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only; requires SC/DPS permit ¦ STATES=Configure/Calibration ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | SetAfePowerState | Method: Set SC-owned AFE global power state with preflight and readback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetAfePowerState | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Slow Controls; DPS may inhibit ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=SC OPC-UA role only ¦ STATES=Off/Standby/Configure/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | ResetAfe | Method: Perform a DAQ configuration reset of an AFE and verify postconditions. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ResetAfe | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ; PDS expert in maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only ¦ STATES=Configure/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | AlignAfes | Method: Run DAQ-owned AFE delay/bitslip alignment and return selected settings. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.AlignAfes | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ; PDS expert in maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only ¦ STATES=Configure/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | SetHdMezzPowerStates | Method: Set paired 5 V and 3.3 V HD-mezzanine rail requests. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetHdMezzPowerStates | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Slow Controls ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=Baseline controlled after commissioning ¦ STATES=Off/Standby/Configure/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | ConfigureHdMezzProtection | Method: Configure shunts/current scales/shutdown limits. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureHdMezzProtection | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=EXCLUDED FROM ROUTINE DCS ¦ AUTHORITY=Slow Controls with PDS expert role ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=SC expert only ¦ STATES=Maintenance only ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | ClearHdMezzAlert | Method: Clear cached/latching HD-mezzanine alert only after the cause is removed. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ClearHdMezzAlert | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=Slow Controls with PDS expert role where required ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=SC OPC-UA role; restricted ¦ STATES=Recovery/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | SetFanOverride | Method: Apply a bounded maintenance fan PWM override; automatic control is default. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetFanOverride | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=EXCLUDED FROM ROUTINE DCS ¦ AUTHORITY=Slow Controls with PDS expert role ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=SC expert only ¦ STATES=Maintenance only ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | ScrapEndpoint | Method: Release DAQ resources, disable data links and restore the agreed non-taking analog profile without changing SC-owned power. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ScrapEndpoint | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only ¦ STATES=Stopped/Configure ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | SoftwareTrigger | Method: Issue a diagnostic software trigger. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SoftwareTrigger | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=CURRENT PROTOBUF ¦ AUTHORITY=DAQ; PDS expert in maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role; not routine DCS ¦ STATES=Calibration/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | SetSpyTriggerControl | Method: Atomically set the spy-buffer source selector and inhibit using inhibit, wait at least five 62.5 MHz clocks, change source, wait again, then release inhibit when requested. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetSpyTriggerControl | String | N/A | Direct | None | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | N/A (write-only method) | Reject without DAQ/expert authority or outside an allowed state; execute inhibit, wait >=5 acquisition clocks, select, wait >=5 clocks, apply the requested final inhibit, and verify both readbacks. | Write-Only | No | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=DAQ; PDS expert in maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role only ¦ STATES=Configure/Calibration/Maintenance ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | DumpSpyBuffers | Method: Request a bounded diagnostic waveform dump. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.DumpSpyBuffers | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=EXCLUDED FROM ROUTINE DCS ¦ AUTHORITY=DAQ; PDS expert in maintenance ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=DAQ OPC-UA role; excluded from routine DCS ¦ STATES=Maintenance/Calibration ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=DAQ ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | DAQ
Operation | WriteAfeRegister | Method: Raw AFE register write. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.WriteAfeRegister | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=EXCLUDED FROM ROUTINE DCS ¦ AUTHORITY=PDS expert ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=Excluded from routine DCS ¦ STATES=Maintenance only ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=PDS expert (maintenance only) ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | PDS expert (maintenance only)
Operation | WriteMultipleFpgaRegisters | Method: Arbitrary FPGA register writes. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.WriteMultipleFpgaRegisters | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=EXCLUDED FROM ROUTINE DCS ¦ AUTHORITY=PDS expert ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=Excluded from routine DCS ¦ STATES=Maintenance only ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=PDS expert (maintenance only) ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | PDS expert (maintenance only)
Operation | FirmwareUpdate | Method: Install/activate a signed DAPHNE firmware release with rollback. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.FirmwareUpdate | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=EXCLUDED FROM ROUTINE DCS ¦ AUTHORITY=Slow Controls with PDS release role ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=SC expert only ¦ STATES=Maintenance only ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls
Operation | RestartBoardService | Method: Restart only the DAPHNE board service and verify a fresh heartbeat. | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.RestartBoardService | String | N/A | Direct | None | [14 empty] | Write-Only | No | On-Change | 1000 | STATUS=REQUIRED PRODUCTION ¦ AUTHORITY=Slow Controls with PDS expert role ¦ INSTANCE=one per DAPHNE board ¦ SOURCE=Workbook-driven OPC-UA policy; gateway backend adapter where mapped ¦ EXPOSURE=SC OPC-UA role; restricted ¦ STATES=Maintenance/Recovery ¦ ALARM=Command audit ¦ INTERLOCK=Guarded ¦ CONTROL_OWNER=Slow Controls ¦ Method result is acceptance/rejection; physical completion requires Operations.LastCommand* and the named readbacks. | Slow Controls

## Registry Metadata
Subsystem | Tag | NodeId pattern | Implementation status | Control owner | Normal authority | Multiplicity | Authoritative/current source | Production exposure | Allowed states | Alarm class | Interlock relevance | DCS access | Poll ms | History rate
Endpoint Status | Connected | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Connected | CURRENT OPC-UA POC | Slow Controls | Read-only fan-out | one per DAPHNE board | NP04 OPC-UA bridge 0.3 / cached DAPHNE poll | Baseline DCS | All states | Stale data | No | Read-Only | 1000 | On-Change
Endpoint Status | StatusSuccess | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Success | CURRENT OPC-UA POC | Slow Controls | Read-only fan-out | one per DAPHNE board | NP04 OPC-UA bridge 0.3 / cached DAPHNE poll | Baseline DCS | All states | Fault | No | Read-Only | 1000 | On-Change
Endpoint Status | StatusMessage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Message | CURRENT OPC-UA POC | Slow Controls | Read-only fan-out | one per DAPHNE board | NP04 OPC-UA bridge 0.3 / cached DAPHNE poll | Baseline DCS | All states | Status | No | Read-Only | 1000 | On-Change
Endpoint Status | LastUpdate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.LastUpdate | CURRENT OPC-UA POC | Slow Controls | Read-only fan-out | one per DAPHNE board | NP04 OPC-UA bridge 0.3 / cached DAPHNE poll | Baseline DCS | All states | Stale data | No | Read-Only | 1000 | On-Change
Endpoint Status | AgeSeconds | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.AgeSeconds | REQUIRED PRODUCTION | Slow Controls | Read-only fan-out | one per DAPHNE board | Bridge freshness contract | Baseline DCS | All states | Stale data | No | Read-Only | 1000 | On-Change
Endpoint Status | Quality | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.Quality | REQUIRED PRODUCTION | Slow Controls | Read-only fan-out | one per DAPHNE board | Bridge quality contract | Baseline DCS | All states | Stale data | No | Read-Only | 1000 | On-Change
Endpoint Status | ErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.ErrorCount | CURRENT OPC-UA POC | Slow Controls | Read-only fan-out | one per DAPHNE board | NP04 OPC-UA bridge ErrorCount | Baseline DCS | All states | Warning | No | Read-Only | 5000 | On-Change
Endpoint Status | ErrorsJsonCompatibility | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Status.ErrorsJson | CURRENT OPC-UA POC | Slow Controls | Read-only fan-out | one per DAPHNE board | NP04 OPC-UA bridge ErrorsJson | Compatibility only | All states | Status | No | Read-Only | 5000 | On-Change
Authority | ActiveWriter | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ActiveWriter | REQUIRED PRODUCTION | Board service (shared audit) | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | AuthorityLeaseId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.AuthorityLeaseId | REQUIRED PRODUCTION | Board service (shared audit) | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | AuthorityLeaseExpiresAt | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.AuthorityLeaseExpiresAt | REQUIRED PRODUCTION | Board service (shared audit) | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | DetectorState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.DetectorState | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | RunNumber | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.RunNumber | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | RunType | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.RunType | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ConfigurationRevision | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ConfigurationRevision | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ConfigurationHash | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ConfigurationHash | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ConfigurationValid | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ConfigurationValid | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | DaqConfigureAllowed | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.DaqConfigureAllowed | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | SlowControlsPermit | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.SlowControlsPermit | REQUIRED PRODUCTION | Slow Controls / DPS | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | DpsPermit | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.DpsPermit | REQUIRED PRODUCTION | Slow Controls / DPS | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ExternalConstraintRevision | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalConstraintRevision | REQUIRED PRODUCTION | Slow Controls / DPS | SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard | one per DAPHNE board | Approved site-specific interaction matrix and SC/DPS interface | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ExternalConstraintStateFresh | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalConstraintStateFresh | REQUIRED PRODUCTION | Slow Controls / DPS | SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard | one per DAPHNE board | Approved site-specific interaction matrix and SC/DPS interface | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ExternalActivityPermit | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalActivityPermit | REQUIRED PRODUCTION | Slow Controls / DPS | SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard | one per DAPHNE board | Approved site-specific interaction matrix and SC/DPS interface | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | ExternalActivityConflict | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.ExternalActivityConflict | REQUIRED PRODUCTION | Slow Controls / DPS | SC/DPS supplies the scoped admission decision; the PDS gateway consumes it read-only and enforces the applicable guard | one per DAPHNE board | Approved site-specific interaction matrix and SC/DPS interface | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | SafeStateActive | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.SafeStateActive | REQUIRED PRODUCTION | Slow Controls / DPS | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | MaintenanceLock | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.MaintenanceLock | REQUIRED PRODUCTION | Slow Controls / DPS | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | LastOvertakeReason | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.LastOvertakeReason | REQUIRED PRODUCTION | Slow Controls / DPS | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Authority | RecoveryRequired | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Authority.RecoveryRequired | REQUIRED PRODUCTION | DAQ | Fixed DAQ/SC field ownership enforced by the OPC-UA gateway and board service; DPS has final inhibit | one per DAPHNE board | Shared OPC-UA gateway role and field-ownership contract | Baseline DCS | All states | State guard | Context | Read-Only | 1000 | On-Change
Command Audit | LastCommandId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandId | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandName | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandName | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandClass | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandClass | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandRequester | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandRequester | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandAuthority | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandAuthority | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandRoute | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandRoute | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandCorrelationId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandCorrelationId | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandRequestTime | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandRequestTime | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandCompletionTime | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandCompletionTime | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandState | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandResult | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandResult | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Command Audit | LastCommandReadbackVerified | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.LastCommandReadbackVerified | REQUIRED PRODUCTION | Board service (shared audit) | Command broker | one per DAPHNE board | Production command-audit contract | Baseline DCS | All states | Audit | No | Read-Only | 1000 | On-Change
Identity | BoardId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.BoardId | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB / deployment inventory | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | HardwareDatabaseId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.HardwareDatabaseId | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | AssetTag | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.AssetTag | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB + physical label | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | Detector | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.Detector | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | DetectorVariant | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.DetectorVariant | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | GeographicLocation | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.GeographicLocation | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | SomProduct | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SomProduct | LOCAL READABLE | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | K26 EEPROM 0x50 | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | SomSerial | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SomSerial | LOCAL READABLE | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | K26 EEPROM 0x50 | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | SomUuid | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SomUuid | LOCAL READABLE | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | K26 EEPROM 0x50 | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | CarrierRevision | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.CarrierRevision | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB + asset record | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | FactoryMac | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.FactoryMac | LOCAL READABLE | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | K26 EEPROM 0x50 | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | ApprovedMac | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.ApprovedMac | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB / network authority | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | MacIdentityMatch | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.MacIdentityMatch | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | Derived identity check | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | Hostname | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.Hostname | LOCAL READABLE | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | Linux + board configuration | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | ExpectedHostname | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.ExpectedHostname | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | HWDB / deployment inventory | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | HostnameMatch | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.HostnameMatch | REQUIRED PRODUCTION | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | Derived identity check | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | TimingEndpointAddress | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.TimingEndpointAddress | REQUIRED PRODUCTION | DAQ / DTS | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | endpointControl.ADDRESS / board inventory | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | CrateId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.CrateId | LOCAL READABLE | DAQ | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | FpgaRegDict.idCrate | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | SlotId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.SlotId | LOCAL READABLE | DAQ | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | FpgaRegDict.idSlot | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | LinkId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.LinkId | LOCAL READABLE | DAQ | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | FpgaRegDict.idLink | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | DetectorId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.DetectorId | LOCAL READABLE | DAQ | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | FpgaRegDict.idDetector | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Identity | RegisterMapVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Identity.RegisterMapVersion | LOCAL READABLE | HWDB / network authority; SC monitors | Read-only; expected values owned by HWDB/network authority | one per DAPHNE board | FpgaRegDict.idVersion | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 60000 | On-Change
Management Network | Name | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Name | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | MacAddress | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.MacAddress | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | PermanentMacAddress | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.PermanentMacAddress | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | Ipv4Address | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Ipv4Address | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | PrefixLength | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.PrefixLength | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | DefaultGateway | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.DefaultGateway | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | DnsServers | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.DnsServers | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | VlanId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.VlanId | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | OperationalState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.OperationalState | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | Carrier | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Carrier | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 10000 | On-Change
Management Network | SpeedMbps | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.SpeedMbps | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | Duplex | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.Duplex | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | MtuBytes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.MtuBytes | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | RxBytes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxBytes | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | RxPackets | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxPackets | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | RxErrors | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxErrors | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | RxDropped | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.RxDropped | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | TxBytes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxBytes | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | TxPackets | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxPackets | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | TxErrors | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxErrors | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | TxDropped | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.TxDropped | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | 10000
Management Network | LinkChangeCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.LinkChangeCount | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 30000 | On-Change
Management Network | ConfigurationMatch | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Network.Interfaces.{Interface}.ConfigurationMatch | REQUIRED PRODUCTION | Slow Controls | Read-only; network assignment owned externally | per interface; baseline instance eth0 | Linux rtnetlink/sysfs plus HWDB/network inventory | Baseline DCS | All states | Network health | No | Read-Only | 10000 | On-Change
Host | CurrentUtc | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.CurrentUtc | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | Linux clock | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | UnixTimeNs | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.UnixTimeNs | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | Linux clock | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | UptimeSeconds | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.UptimeSeconds | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | Linux /proc/uptime | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | BootId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.BootId | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | Linux /proc/sys/kernel/random/boot_id | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | BootTime | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.BootTime | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | CurrentUtc - UptimeSeconds | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | KernelRelease | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.KernelRelease | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | uname | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | OperatingSystemVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.OperatingSystemVersion | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | /etc/os-release | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | RootFilesystemBuildId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.RootFilesystemBuildId | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | release manifest | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | ActiveBootSlot | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.ActiveBootSlot | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | boot-state service/U-Boot env | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | BootMarkedGood | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.BootMarkedGood | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | daphne-boot-ok.service | Baseline DCS | All states | Host health | No | Read-Only | 10000 | On-Change
Host | LastResetReason | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.LastResetReason | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | kernel/PMU/boot evidence | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | WatchdogHealthy | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.WatchdogHealthy | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | watchdog/sysfs | Baseline DCS | All states | Host health | No | Read-Only | 10000 | On-Change
Host | NtpSynchronized | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.NtpSynchronized | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | timesync service | Baseline DCS | All states | Host health | No | Read-Only | 10000 | On-Change
Host | NtpOffsetMilliseconds | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.NtpOffsetMilliseconds | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | timesync service | Baseline DCS | All states | Host health | No | Read-Only | 10000 | On-Change
Host | TimeSource | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.TimeSource | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | timesync service | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | CpuLoad1Minute | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.CpuLoad1Minute | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | /proc/loadavg | Baseline DCS | All states | Host health | No | Read-Only | 10000 | On-Change
Host | MemoryAvailableBytes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.MemoryAvailableBytes | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | /proc/meminfo | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | RootFilesystemFreeBytes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.RootFilesystemFreeBytes | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | statvfs | Baseline DCS | All states | Host health | No | Read-Only | 60000 | On-Change
Host | RootFilesystemReadOnly | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Host.RootFilesystemReadOnly | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | mount state | Baseline DCS | All states | Host health | No | Read-Only | 10000 | On-Change
Services | Active | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.Active | CURRENT PROTOBUF | Slow Controls | Read-only; service recovery owned by PDS operations | per required service; see Instance Sets | SlowControlStatusResponse.ServiceStatus / systemd | Baseline DCS | All states | Service fault | Readiness | Read-Only | 5000 | On-Change
Services | State | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.State | CURRENT PROTOBUF | Slow Controls | Read-only; service recovery owned by PDS operations | per required service; see Instance Sets | SlowControlStatusResponse.ServiceStatus / systemd | Baseline DCS | All states | Service fault | Readiness | Read-Only | 5000 | On-Change
Services | MainPid | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.MainPid | REQUIRED PRODUCTION | Slow Controls | Read-only; service recovery owned by PDS operations | per required service; see Instance Sets | SlowControlStatusResponse.ServiceStatus / systemd | Baseline DCS | All states | Service fault | Readiness | Read-Only | 5000 | On-Change
Services | RestartCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.RestartCount | REQUIRED PRODUCTION | Slow Controls | Read-only; service recovery owned by PDS operations | per required service; see Instance Sets | SlowControlStatusResponse.ServiceStatus / systemd | Baseline DCS | All states | Service fault | Readiness | Read-Only | 5000 | On-Change
Services | StartTime | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.StartTime | REQUIRED PRODUCTION | Slow Controls | Read-only; service recovery owned by PDS operations | per required service; see Instance Sets | SlowControlStatusResponse.ServiceStatus / systemd | Baseline DCS | All states | Service fault | Readiness | Read-Only | 5000 | On-Change
Services | LastExitStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Services.{Service}.LastExitStatus | REQUIRED PRODUCTION | Slow Controls | Read-only; service recovery owned by PDS operations | per required service; see Instance Sets | SlowControlStatusResponse.ServiceStatus / systemd | Baseline DCS | All states | Service fault | Readiness | Read-Only | 5000 | On-Change
OPC-UA Bridge | Heartbeat | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.Heartbeat | CURRENT OPC-UA POC | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 1000 | On-Change
OPC-UA Bridge | Version | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.Version | CURRENT OPC-UA POC | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | ConfigurationPath | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ConfigurationPath | CURRENT OPC-UA POC | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | ConfigurationHash | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ConfigurationHash | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | NamespaceVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.NamespaceVersion | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | BoardCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.BoardCount | CURRENT OPC-UA POC | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | WorkerCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.WorkerCount | CURRENT OPC-UA POC | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | CertificateNotAfter | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.CertificateNotAfter | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | RejectedSessionCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.RejectedSessionCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | PollTimeoutCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.PollTimeoutCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | PollErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.PollErrorCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | BackendProtocolVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.BackendProtocolVersion | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | BackendCompatibilityState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.BackendCompatibilityState | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | ActiveDaqSessionCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ActiveDaqSessionCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | ActiveSlowControlsSessionCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.ActiveSlowControlsSessionCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | UnauthorizedMethodCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.UnauthorizedMethodCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | CommandQueueDepth | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.CommandQueueDepth | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
OPC-UA Bridge | TranslationErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Bridge.TranslationErrorCount | REQUIRED PRODUCTION | PDS OPC-UA gateway service | Neutral PDS gateway deployment shared by DAQ and SC | one per bridge shard | NP04 bridge for current fields; production shared-gateway contract otherwise | Baseline DCS | All states | Bridge health | No | Read-Only | 30000 | On-Change
Firmware | Loaded | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.Loaded | CURRENT PROTOBUF | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | SlowControlStatusResponse.FirmwareStatus.loaded | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | FpgaManagerState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FpgaManagerState | CURRENT PROTOBUF | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | FirmwareStatus.fpga_manager_state | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | OverlayName | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.OverlayName | CURRENT PROTOBUF | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | FirmwareStatus.overlay_name | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | BuildId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.BuildId | CURRENT PROTOBUF | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | FirmwareStatus.build_id | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | GitCommit | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GitCommit | LOCAL READABLE | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | FpgaRegDict.GIT | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | ExpectedBuildId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ExpectedBuildId | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | release manifest | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | BuildMatchesExpected | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.BuildMatchesExpected | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | derived comparison | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | RegisterMapVersionMatches | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.RegisterMapVersionMatches | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | idVersion + server manifest | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | PlI2cDevicePresent | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.PlI2cDevicePresent | CURRENT PROTOBUF | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | FirmwareStatus.pl_devices_present / preflight | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | PlSpiDevicePresent | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.PlSpiDevicePresent | CURRENT PROTOBUF | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | FirmwareStatus.pl_devices_present / preflight | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | ActiveXmutilApplication | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ActiveXmutilApplication | LOCAL READABLE | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | xmutil listapps | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | FirmwareUpdateState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FirmwareUpdateState | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | firmware update contract | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | FirmwarePathAvailable | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FirmwarePathAvailable | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | firmware update contract | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | FirmwareArtifactSha256 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.FirmwareArtifactSha256 | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | release/deployment manifest | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | ServerVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ServerVersion | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | board service build metadata | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | ProtobufSchemaVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ProtobufSchemaVersion | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | board service build metadata | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | GatewareHeartbeat | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GatewareHeartbeat | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | required gateware health block | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | GatewareFatalFaultLatched | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GatewareFatalFaultLatched | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | required gateware health block | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
Firmware | GatewareResetReason | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.GatewareResetReason | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | required gateware health block | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 30000 | On-Change
Firmware | ConfigurationReady | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Firmware.ConfigurationReady | REQUIRED PRODUCTION | Slow Controls | Read-only; updates are expert maintenance | one per DAPHNE board | analog-control contract | Baseline DCS | All states | Firmware readiness | Readiness | Read-Only | 5000 | On-Change
RPU | Available | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Available | CURRENT PROTOBUF | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Context | Read-Only | 1000 | On-Change
RPU | Running | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Running | CURRENT PROTOBUF | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Yes | Read-Only | 1000 | On-Change
RPU | Firmware | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Firmware | CURRENT PROTOBUF | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Context | Read-Only | 1000 | On-Change
RPU | Heartbeat | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.Heartbeat | CURRENT PROTOBUF | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Yes | Read-Only | 1000 | On-Change
RPU | LastFault | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.LastFault | CURRENT PROTOBUF | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Context | Read-Only | 1000 | On-Change
RPU | AbiVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.AbiVersion | REQUIRED PRODUCTION | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Context | Read-Only | 1000 | On-Change
RPU | InterlockReady | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.InterlockReady | REQUIRED PRODUCTION | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Yes | Read-Only | 1000 | On-Change
RPU | FaultCode | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.FaultCode | REQUIRED PRODUCTION | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Yes | Read-Only | 1000 | On-Change
RPU | CommandQueueDepth | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.CommandQueueDepth | REQUIRED PRODUCTION | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Context | Read-Only | 1000 | On-Change
RPU | LastReplySequence | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.RPU.LastReplySequence | REQUIRED PRODUCTION | Slow Controls | Read-only; RPU owns safety-critical AFE sequencing | one per DAPHNE board | SlowControlStatusResponse.RpuStatus / RPU wire status | Baseline DCS | All states | Control-path fault | Context | Read-Only | 1000 | On-Change
I2C Bus | Present | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Present | CURRENT PROTOBUF | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | Path | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Path | CURRENT PROTOBUF | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | Functional | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Functional | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | Owner | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Owner | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | Busy | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Busy | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | LastSuccess | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.LastSuccess | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | ErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.ErrorCount | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | TimeoutCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.TimeoutCount | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Bus | RecoveryCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.RecoveryCount | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized hardware owner | per I2C bus; expected buses 1 and 2 | SlowControlStatusResponse.I2CBusStatus plus required metrics | Baseline DCS | All states | I2C health | Readiness | Read-Only | 5000 | On-Change
I2C Device | Name | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Name | CURRENT PROTOBUF | Slow Controls | Read-only | per expected and discovered I2C device; see Instance Sets | SlowControlStatusResponse.I2CDeviceStatus / device driver | Baseline DCS | All states | I2C device fault | Readiness | Read-Only | 10000 | On-Change
I2C Device | Present | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Present | CURRENT PROTOBUF | Slow Controls | Read-only | per expected and discovered I2C device; see Instance Sets | SlowControlStatusResponse.I2CDeviceStatus / device driver | Baseline DCS | All states | I2C device fault | Readiness | Read-Only | 10000 | On-Change
I2C Device | Driver | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Driver | REQUIRED PRODUCTION | Slow Controls | Read-only | per expected and discovered I2C device; see Instance Sets | SlowControlStatusResponse.I2CDeviceStatus / device driver | Baseline DCS | All states | I2C device fault | Readiness | Read-Only | 10000 | On-Change
I2C Device | Status | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.Status | CURRENT PROTOBUF | Slow Controls | Read-only | per expected and discovered I2C device; see Instance Sets | SlowControlStatusResponse.I2CDeviceStatus / device driver | Baseline DCS | All states | I2C device fault | Readiness | Read-Only | 10000 | On-Change
I2C Device | LastSuccess | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.LastSuccess | REQUIRED PRODUCTION | Slow Controls | Read-only | per expected and discovered I2C device; see Instance Sets | SlowControlStatusResponse.I2CDeviceStatus / device driver | Baseline DCS | All states | I2C device fault | Readiness | Read-Only | 10000 | On-Change
I2C Device | ErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.I2C.Buses.{Bus}.Devices.{Address}.ErrorCount | REQUIRED PRODUCTION | Slow Controls | Read-only | per expected and discovered I2C device; see Instance Sets | SlowControlStatusResponse.I2CDeviceStatus / device driver | Baseline DCS | All states | I2C device fault | Readiness | Read-Only | 10000 | On-Change
SPI | Present | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.Present | LOCAL READABLE | Slow Controls | Read-only; one serialized owner | per SPI device; baseline current monitor | preflight /dev/spidev3.0 and current-monitor driver | Baseline DCS | All states | SPI health | Readiness | Read-Only | 5000 | On-Change
SPI | Functional | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.Functional | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized owner | per SPI device; baseline current monitor | preflight /dev/spidev3.0 and current-monitor driver | Baseline DCS | All states | SPI health | Readiness | Read-Only | 5000 | On-Change
SPI | Owner | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.Owner | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized owner | per SPI device; baseline current monitor | preflight /dev/spidev3.0 and current-monitor driver | Baseline DCS | All states | SPI health | Readiness | Read-Only | 5000 | On-Change
SPI | ErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.ErrorCount | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized owner | per SPI device; baseline current monitor | preflight /dev/spidev3.0 and current-monitor driver | Baseline DCS | All states | SPI health | Readiness | Read-Only | 5000 | On-Change
SPI | LastSuccess | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.SPI.{Device}.LastSuccess | REQUIRED PRODUCTION | Slow Controls | Read-only; one serialized owner | per SPI device; baseline current monitor | preflight /dev/spidev3.0 and current-monitor driver | Baseline DCS | All states | SPI health | Readiness | Read-Only | 5000 | On-Change
Board Rails | 3V3PDSVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3PDS.Voltage | CURRENT PROTOBUF | Slow Controls | Read-only | one per DAPHNE board | ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | 1000
Board Rails | 3V3PDSStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3PDS.Status | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | required typed rail status | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | On-Change
Board Rails | 1V8PDSVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8PDS.Voltage | LOCAL READABLE | Slow Controls | Read-only | one per DAPHNE board | ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | 1000
Board Rails | 1V8PDSStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8PDS.Status | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | required typed rail status | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | On-Change
Board Rails | 3V3AVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3A.Voltage | LOCAL READABLE | Slow Controls | Read-only | one per DAPHNE board | ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | 1000
Board Rails | 3V3AStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.3V3A.Status | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | required typed rail status | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | On-Change
Board Rails | 1V8AVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8A.Voltage | CURRENT PROTOBUF | Slow Controls | Read-only | one per DAPHNE board | ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | 1000
Board Rails | 1V8AStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.1V8A.Status | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | required typed rail status | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | On-Change
Board Rails | Minus5VAVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.Minus5VA.Voltage | CURRENT PROTOBUF | Slow Controls | Read-only | one per DAPHNE board | ADS7138 0x10/0x17 cached monitor; GeneralInfo exposes a subset | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | 1000
Board Rails | Minus5VAStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.BoardRails.Minus5VA.Status | REQUIRED PRODUCTION | Slow Controls | Read-only | one per DAPHNE board | required typed rail status | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 1000 | On-Change
PMBus Regulators | Voltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Voltage | CURRENT PROTOBUF | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | 5000
PMBus Regulators | Current | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Current | CURRENT PROTOBUF | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | 5000
PMBus Regulators | Power | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Power | CURRENT PROTOBUF | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | 5000
PMBus Regulators | Temperature | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Temperature | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | 5000
PMBus Regulators | StatusByte | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusByte | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
PMBus Regulators | StatusWord | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusWord | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
PMBus Regulators | StatusVout | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusVout | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
PMBus Regulators | StatusIout | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusIout | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
PMBus Regulators | StatusTemperature | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusTemperature | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
PMBus Regulators | StatusCml | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.StatusCml | LOCAL READABLE | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
PMBus Regulators | Health | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Power.PMBus.{Rail}.Health | REQUIRED PRODUCTION | Slow Controls | Read-only | per PMBus regulator: 3VD3, 2VA1, 3VA6, 1VD8 | PJT004A0X43 / Rust PMBus collector | Baseline DCS | All states | Rail fault | Candidate | Read-Only | 5000 | On-Change
Temperature | Celsius | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Sensors.{Sensor}.Celsius | CURRENT PROTOBUF | Slow Controls | Read-only | per discovered allow-listed sensor | SlowControlStatusResponse.TemperatureReading; Linux thermal/hwmon; PMBus/MCP9808/SFP | Baseline DCS | All states | Thermal fault | Candidate | Read-Only | 5000 | 5000
Temperature | Status | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Sensors.{Sensor}.Status | REQUIRED PRODUCTION | Slow Controls | Read-only | per discovered allow-listed sensor | SlowControlStatusResponse.TemperatureReading; Linux thermal/hwmon; PMBus/MCP9808/SFP | Baseline DCS | All states | Thermal fault | Candidate | Read-Only | 5000 | On-Change
Temperature | LastUpdate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Sensors.{Sensor}.LastUpdate | REQUIRED PRODUCTION | Slow Controls | Read-only | per discovered allow-listed sensor | SlowControlStatusResponse.TemperatureReading; Linux thermal/hwmon; PMBus/MCP9808/SFP | Baseline DCS | All states | Thermal fault | Candidate | Read-Only | 5000 | On-Change
HD Mezzanine | BlockEnabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.BlockEnabled | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | Configured | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Configured | REQUIRED PRODUCTION | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | Power5VRequested | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power5VRequested | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | Power3V3Requested | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power3V3Requested | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | Voltage5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Voltage5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | Current5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Current5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | Power5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | Alert5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Alert5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | Voltage3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Voltage3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | Current3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Current3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | Power3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Power3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | Alert3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.Alert3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | RShunt5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.RShunt5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | RShunt3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.RShunt3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | MaxCurrentScale5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxCurrentScale5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | MaxCurrentScale3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxCurrentScale3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | ShutdownCurrent5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShutdownCurrent5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | ShutdownCurrent3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShutdownCurrent3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | MaxPower5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxPower5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | MaxPower3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.MaxPower3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | CurrentLsb5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.CurrentLsb5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | CurrentLsb3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.CurrentLsb3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | 1000
HD Mezzanine | ShuntCal5V | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShuntCal5V | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
HD Mezzanine | ShuntCal3V3 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.HDMezz.{Afe}.ShuntCal3V3 | CURRENT PROTOBUF | Slow Controls | SC owns rail requests; monitor values are read-only | per AFE block 0..4 when HD mezzanine is installed | HD mezzanine Protobuf API and cached INA232/TCA9536 driver | Baseline DCS | All states | Mezzanine power fault | Candidate | Read-Only | 1000 | On-Change
SFP | Present | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.Present | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 5000 | On-Change
SFP | LossOfSignal | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.LossOfSignal | LOCAL READABLE | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 5000 | On-Change
SFP | TxFault | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxFault | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 5000 | On-Change
SFP | TxDisabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxDisabled | LOCAL READABLE | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 5000 | On-Change
SFP | RateSelect | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.RateSelect | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | VendorName | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.VendorName | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | VendorOui | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.VendorOui | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | PartNumber | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.PartNumber | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | Revision | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.Revision | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | SerialNumber | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.SerialNumber | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | DateCode | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.DateCode | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | NominalBitRateMbps | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.NominalBitRateMbps | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | WavelengthNm | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.WavelengthNm | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFF-8472 plus EDMS 2580777 endpoint SFP specification | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | 5000
SFP | TemperatureCelsius | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TemperatureCelsius | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | 5000
SFP | SupplyVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.SupplyVoltage | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | 5000
SFP | TxBiasCurrent | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxBiasCurrent | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | 5000
SFP | TxOpticalPower | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.TxOpticalPower | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | 5000
SFP | RxOpticalPower | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.RxOpticalPower | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | 5000
SFP | AlarmFlags | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.AlarmFlags | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | WarningFlags | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.WarningFlags | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 30000 | On-Change
SFP | DomSupported | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.DomSupported | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 5000 | On-Change
SFP | EepromReadable | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Optics.SFP.{Sfp}.EepromReadable | REQUIRED PRODUCTION | Slow Controls | Read-only; TX disable controlled by endpoint/link firmware | per timing, CCM, and provisioned data SFP | SFP GPIO/I2C expander 0x72 and SFF-8472 | Baseline DCS | All states | Optical link fault | Readiness | Read-Only | 5000 | On-Change
Fans | Present | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.Present | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | On-Change
Fans | PwmCommand | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.PwmCommand | LOCAL READABLE | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | 1000
Fans | TachometerRaw | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.TachometerRaw | LOCAL READABLE | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | 1000
Fans | SpeedRpm | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.SpeedRpm | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | 1000
Fans | Running | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.Running | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | On-Change
Fans | Stalled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.Stalled | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | On-Change
Fans | StallCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.StallCount | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | On-Change
Fans | ControlMode | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.ControlMode | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | On-Change
Fans | LastUpdate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Thermal.Fans.{Fan}.LastUpdate | REQUIRED PRODUCTION | Slow Controls | Local automatic control; maintenance override only | per fan 0..1 | FpgaRegDict fanControl/fanReadSpeed_0..1 | Baseline DCS | All states | Fan fault | Candidate | Read-Only | 1000 | On-Change
Timing | ClockSourceControlled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.ClockSourceControlled | CURRENT PROTOBUF | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | endpointClockControl.CLOCK_SOURCE | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | Mmcm0Locked | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.Mmcm0Locked | CURRENT OPC-UA POC | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | ClockStatus.mmcm0_locked | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | Mmcm1Locked | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.Mmcm1Locked | CURRENT OPC-UA POC | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | ClockStatus.mmcm1_locked | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | EndpointFsmState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointFsmState | LOCAL READABLE | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | endpointStatus.FSM_STATUS | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | EndpointFsmRaw | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointFsmRaw | LOCAL READABLE | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | endpointStatus.FSM_STATUS | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | TimestampValid | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TimestampValid | LOCAL READABLE | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | endpointStatus.TIMESTAMP_OK | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | TimingUsable | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TimingUsable | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | preflight readiness contract | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | RawEndpointStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.RawEndpointStatus | CURRENT PROTOBUF | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | ClockStatus.raw_endpoint_status | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | EndpointAddress | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointAddress | LOCAL READABLE | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | endpointControl.ADDRESS | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | TxEnableRequested | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TxEnableRequested | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | timing endpoint contract | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | TxDisabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.TxDisabled | LOCAL READABLE | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | timing endpoint output | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | LossOfSignal | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LossOfSignal | LOCAL READABLE | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | sfp_tmg_los | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | LockLossCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LockLossCount | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | required timing health block | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | AutoRelockCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.AutoRelockCount | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | required timing health block | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | LastLockLossTime | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LastLockLossTime | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | required timing health block | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | Timestamp | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.Timestamp | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | timestamp0..3 register words | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | LastDtsCommandId | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.LastDtsCommandId | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | timing/header integration contract | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | GroupMask | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.GroupMask | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | versioned timing configuration | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | DelayTrim | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.DelayTrim | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | versioned timing configuration | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | CommandDecoderMap | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.CommandDecoderMap | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | EDMS 2395364 v2 control-bus registers 0x60..0x67 | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | ClockPhaseUpdateComplete | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.ClockPhaseUpdateComplete | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | EDMS 2395364 v2 status register 0x71 bit 4 | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | CoarseDelay | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.CoarseDelay | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | EDMS 2395364 v2 deskew registers 0x72..0x73 bits 3..0 | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | FinePhase | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.FinePhase | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | EDMS 2395364 v2 deskew registers 0x72..0x73 bits 15..4 | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | ProtocolErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.ProtocolErrorCount | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | EDMS 2395364 v2 err_cnt register 0x76 | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
Timing | EndpointVersion | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Timing.EndpointVersion | REQUIRED PRODUCTION | DAQ / DTS | DAQ/DTS owns timing configuration; SC receives monitoring; DPS may inhibit | one per DAPHNE board | EDMS 2395364 v2 version registers 0x78..0x7b | Baseline DCS | All states | Timing readiness | Readiness | Read-Only | 1000 | On-Change
AFE Global | PowerState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.PowerState | LOCAL READABLE | Slow Controls | SC owns power/bias enable and safe-state action | one per DAPHNE board | afeGlobalControl.POWERSTATE | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE Global | ResetAsserted | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.ResetAsserted | LOCAL READABLE | DAQ | DAQ owns data-taking configuration/reset; SC monitors | one per DAPHNE board | afeGlobalControl.RESET | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE Global | BusyAfe0 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BusyAfe0 | LOCAL READABLE | DAQ | DAQ owns data-taking configuration/reset; SC monitors | one per DAPHNE board | afeGlobalControl.BUSY_AFE0 | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE Global | BusyAfe12 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BusyAfe12 | LOCAL READABLE | DAQ | DAQ owns data-taking configuration/reset; SC monitors | one per DAPHNE board | afeGlobalControl.BUSY_AFE12 | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE Global | BusyAfe34 | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BusyAfe34 | LOCAL READABLE | DAQ | DAQ owns data-taking configuration/reset; SC monitors | one per DAPHNE board | afeGlobalControl.BUSY_AFE34 | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE Global | BiasEnable | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BiasEnable | LOCAL READABLE | Slow Controls | SC owns power/bias enable and safe-state action | one per DAPHNE board | biasEnable.ENABLE | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE Global | BiasControlCode | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Global.BiasControlCode | CURRENT PROTOBUF | DAQ | DAQ owns data-taking configuration/reset; SC monitors | one per DAPHNE board | READ_VBIAS_CONTROL | Baseline DCS | All states | AFE readiness | Candidate | Read-Only | 1000 | On-Change
AFE | Attenuation | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.Attenuation | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | READ/WRITE AFE attenuation | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | VGain | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.VGain | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | READ_AFE_VGAIN | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | BiasSetCode | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasSetCode | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | READ_AFE_BIAS_SET | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | BiasVoltage | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltage | CURRENT OPC-UA POC | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | GeneralInfo.v_bias_0..4 / bias monitor | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | 5000
AFE | BiasVoltageTarget | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltageTarget | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | cmd_writeAFEBiasControlledSet.target_bias_mv | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | 5000
AFE | BiasVoltageError | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltageError | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | cmd_writeAFEBiasControlledSet_response.error_mv | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | 5000
AFE | BiasControlStatus | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasControlStatus | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AfeBiasControlStatus | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | BiasControlSettled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasControlSettled | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | controlled-set response.settled | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | BiasControlIterations | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasControlIterations | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | controlled-set response.iterations | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | BiasVoltageSlope | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasVoltageSlope | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | controlled-set response.final_slope_mv_per_s | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | 5000
AFE | BiasStableSamples | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.BiasStableSamples | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | controlled-set response.stable_samples | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | AdcResolution | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AdcResolution | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.adc.resolution | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | AdcOutputFormat | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AdcOutputFormat | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.adc.output_format | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | AdcMsbFirst | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AdcMsbFirst | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.adc.sb_first | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | PgaLpfCutFrequency | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.PgaLpfCutFrequency | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.pga.lpf_cut_frequency | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | PgaIntegratorDisabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.PgaIntegratorDisabled | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.pga.integrator_disable | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | PgaGain | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.PgaGain | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.pga.gain | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | LnaClamp | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.LnaClamp | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.lna.clamp | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | LnaGain | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.LnaGain | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.lna.gain | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | LnaIntegratorDisabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.LnaIntegratorDisabled | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | AFEConfig.lna.integrator_disable | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | AlignmentDelay | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AlignmentDelay | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | cmd_alignAFEs_response.delay | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | AlignmentBitslip | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.AlignmentBitslip | CURRENT PROTOBUF | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | cmd_alignAFEs_response.bitslip | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
AFE | Aligned | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.AFE.Blocks.{Afe}.Aligned | REQUIRED PRODUCTION | DAQ | DAQ applies the data-taking profile through its OPC-UA role; SC monitors and owns energization/permits | per AFE block 0..4 | frontend alignment status contract | Baseline DCS | All states | AFE configuration/readback | Context | Read-Only | 5000 | On-Change
Channels | Trim | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Trim | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | READ_TRIM_* | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | Offset | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Offset | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | READ_OFFSET_* | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | Gain | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Gain | REQUIRED PRODUCTION | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | ChannelConfig.gain + readback contract | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | CurrentMonitorRaw | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.CurrentMonitorRaw | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | READ_CURRENT_MONITOR | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | Current | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.Current | REQUIRED PRODUCTION | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | current monitor + calibration | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | TriggerEnabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerEnabled | LOCAL READABLE | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | triggerEnableLow/High | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | FullStreamEnabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.FullStreamEnabled | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | ConfigureRequest.full_stream_channels | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | TriggerThreshold | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerThreshold | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | TriggerChannelSnapshot.threshold | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | TriggerRecordCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerRecordCount | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | TriggerChannelSnapshot.record_count | Baseline DCS | All states | Run context | No | Read-Only | 30000 | 30000
Channels | TriggerBusyCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerBusyCount | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | TriggerChannelSnapshot.busy_count | Baseline DCS | All states | Run context | No | Read-Only | 30000 | 30000
Channels | TriggerFullCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerFullCount | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | TriggerChannelSnapshot.full_count | Baseline DCS | All states | Run context | No | Read-Only | 30000 | 30000
Channels | TriggerRate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.TriggerRate | REQUIRED PRODUCTION | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | delta(record_count)/delta(time) | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | BusyFraction | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.BusyFraction | REQUIRED PRODUCTION | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | delta(busy)/defined opportunity denominator | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Channels | PacketCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.PacketCount | CURRENT PROTOBUF | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | ChannelInfo.total_packets | Baseline DCS | All states | Run context | No | Read-Only | 30000 | 30000
Channels | PacketRate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Channels.{Channel}.PacketRate | REQUIRED PRODUCTION | DAQ | DAQ applies channel configuration through its OPC-UA role; SC receives the complete readback/monitoring set | per channel; HD 0..39, VD 0..31 | ChannelInfo.new_packets / interval | Baseline DCS | All states | Run context | No | Read-Only | 5000 | 5000
Run Configuration | SelfTriggerThreshold | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerThreshold | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.self_trigger_threshold | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | SelfTriggerXcorr | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerXcorr | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.self_trigger_xcorr | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | TriggerPrimitiveConfig | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.TriggerPrimitiveConfig | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.tp_conf | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | CompensatorMask | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.CompensatorMask | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.compensator | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | InverterMask | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.InverterMask | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.inverters | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | FullStreamChannelList | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.FullStreamChannelList | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.full_stream_channels | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | Slot | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.Slot | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.slot | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | RequestTimeoutMs | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.RequestTimeoutMs | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.timeout_ms | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | BiasControl | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.BiasControl | CURRENT PROTOBUF | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | ConfigureRequest.biasctrl | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | SelfTriggerConfigRaw | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerConfigRaw | LOCAL READABLE | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | stuffregs.OFF_ST_CONFIG | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | SelfTriggerDelayRaw | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerDelayRaw | LOCAL READABLE | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | stuffregs.OFF_ST_DELAY | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Run Configuration | SelfTriggerFilterOutputSelect | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Configuration.Active.SelfTriggerFilterOutputSelect | LOCAL READABLE | DAQ | DAQ applies and verifies the run configuration through the OPC-UA gateway; SC monitors readback | one per DAPHNE board | stuffregs.OFF_ST_FILTER_OUTPUT_SEL | Baseline DCS | All states | Configuration mismatch | No | Read-Only | 30000 | On-Change
Spy Trigger | SourceSelector | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Spy.Trigger.SourceSelector | LOCAL READABLE | DAQ | DAQ OPC-UA configuration; SC monitoring | one per DAPHNE board | daphne-firmware working tree: spy_trigger_ctrl_reg[1:0] at 0x88000034 | Baseline DCS readback; write only through guarded typed operation | All states | Spy-buffer configuration | Context | Read-Only | 5000 | On-Change
Spy Trigger | Inhibit | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Spy.Trigger.Inhibit | LOCAL READABLE | DAQ | DAQ OPC-UA configuration; SC monitoring | one per DAPHNE board | daphne-firmware working tree: spy_trigger_ctrl_reg[2] at 0x88000034 | Baseline DCS readback; write only through guarded typed operation | All states | Spy capture inhibited | Context | Read-Only | 5000 | On-Change
Readout Stream | TotalPackets | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.TotalPackets | CURRENT PROTOBUF | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | NewPackets | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.NewPackets | CURRENT PROTOBUF | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | TotalDroppedPackets | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.TotalDroppedPackets | CURRENT PROTOBUF | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | NewDroppedPackets | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.NewDroppedPackets | CURRENT PROTOBUF | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | PacketRate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.PacketRate | REQUIRED PRODUCTION | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | DropRate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.DropRate | REQUIRED PRODUCTION | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | DropFraction | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.DropFraction | REQUIRED PRODUCTION | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | ThrottlingActive | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.ThrottlingActive | REQUIRED PRODUCTION | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | ThrottleEventCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.ThrottleEventCount | REQUIRED PRODUCTION | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
Readout Stream | DataPathStale | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Stream.DataPathStale | REQUIRED PRODUCTION | DAQ | Read-only; DAQ owns data-path recovery | one per DAPHNE board | InfoResponse.StreamInfo plus required gateware/Hermes health | Baseline DCS | All states | Readout readiness | No | Read-Only | 5000 | 5000
10GbE Readout | Provisioned | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.Provisioned | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | Enabled | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.Enabled | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | LinkReady | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.LinkReady | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | LocalFault | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.LocalFault | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | RemoteFault | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.RemoteFault | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | TxPacketCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.TxPacketCount | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | TxByteCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.TxByteCount | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | ErrorCount | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.ErrorCount | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | DestinationIp | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.DestinationIp | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | DestinationPort | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.DestinationPort | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
10GbE Readout | MtuBytes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Readout.Links.{DataLink}.MtuBytes | REQUIRED PRODUCTION | DAQ | DAQ supplies the readout-link profile through the OPC-UA gateway and monitors the data plane; SC receives health readback | per provisioned 10 GbE data lane; hardware supports up to four | Hermes/10GbE gateware health contract | Baseline DCS | All states | Readout link fault | No | Read-Only | 5000 | On-Change
Operation | RequestSafeState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.RequestSafeState | REQUIRED PRODUCTION | Slow Controls / DPS | Slow Controls/DPS | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | Baseline controlled | Normal, Running, Fault | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | AcknowledgeRecovery | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.AcknowledgeRecovery | REQUIRED PRODUCTION | Slow Controls | Slow Controls | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | Baseline controlled | Recovery | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ConfigureRun | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureRun | CURRENT PROTOBUF | DAQ | DAQ | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only; not an Ignition operator button | Configure/Calibration | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ConfigureClocks | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureClocks | CURRENT PROTOBUF | DAQ / DTS | DAQ/DTS; PDS expert in maintenance | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ/DTS OPC-UA roles only | Configure/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ConfigureFrontend | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureFrontend | CURRENT PROTOBUF | DAQ | DAQ | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only; not a casual operator write | Configure/Calibration | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | SetControlledBias | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetControlledBias | CURRENT PROTOBUF | DAQ | DAQ | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only; requires SC/DPS permit | Configure/Calibration | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | SetAfePowerState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetAfePowerState | CURRENT PROTOBUF | Slow Controls | Slow Controls; DPS may inhibit | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | SC OPC-UA role only | Off/Standby/Configure/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ResetAfe | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ResetAfe | CURRENT PROTOBUF | DAQ | DAQ; PDS expert in maintenance | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only | Configure/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | AlignAfes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.AlignAfes | CURRENT PROTOBUF | DAQ | DAQ; PDS expert in maintenance | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only | Configure/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | SetHdMezzPowerStates | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetHdMezzPowerStates | CURRENT PROTOBUF | Slow Controls | Slow Controls | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | Baseline controlled after commissioning | Off/Standby/Configure/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ConfigureHdMezzProtection | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureHdMezzProtection | EXCLUDED FROM ROUTINE DCS | Slow Controls | Slow Controls with PDS expert role | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | SC expert only | Maintenance only | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ClearHdMezzAlert | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ClearHdMezzAlert | CURRENT PROTOBUF | Slow Controls | Slow Controls with PDS expert role where required | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | SC OPC-UA role; restricted | Recovery/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | SetFanOverride | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetFanOverride | EXCLUDED FROM ROUTINE DCS | Slow Controls | Slow Controls with PDS expert role | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | SC expert only | Maintenance only | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | ScrapEndpoint | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ScrapEndpoint | CURRENT PROTOBUF | DAQ | DAQ | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only | Stopped/Configure | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | SoftwareTrigger | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SoftwareTrigger | CURRENT PROTOBUF | DAQ | DAQ; PDS expert in maintenance | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role; not routine DCS | Calibration/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | SetSpyTriggerControl | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetSpyTriggerControl | REQUIRED PRODUCTION | DAQ | DAQ; PDS expert in maintenance | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role only | Configure/Calibration/Maintenance | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | DumpSpyBuffers | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.DumpSpyBuffers | EXCLUDED FROM ROUTINE DCS | DAQ | DAQ; PDS expert in maintenance | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | DAQ OPC-UA role; excluded from routine DCS | Maintenance/Calibration | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | WriteAfeRegister | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.WriteAfeRegister | EXCLUDED FROM ROUTINE DCS | PDS expert (maintenance only) | PDS expert | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | Excluded from routine DCS | Maintenance only | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | WriteMultipleFpgaRegisters | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.WriteMultipleFpgaRegisters | EXCLUDED FROM ROUTINE DCS | PDS expert (maintenance only) | PDS expert | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | Excluded from routine DCS | Maintenance only | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | FirmwareUpdate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.FirmwareUpdate | EXCLUDED FROM ROUTINE DCS | Slow Controls | Slow Controls with PDS release role | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | SC expert only | Maintenance only | Command audit | Guarded | Write-Only | 1000 | On-Change
Operation | RestartBoardService | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.RestartBoardService | REQUIRED PRODUCTION | Slow Controls | Slow Controls with PDS expert role | one per DAPHNE board | Workbook-driven OPC-UA policy; gateway backend adapter where mapped | SC OPC-UA role; restricted | Maintenance/Recovery | Command audit | Guarded | Write-Only | 1000 | On-Change

## Instance Sets
Placeholder | Entity | Allowed/expected instances | Source of truth
{BoardId} | DAPHNE board | One instance per installed board from HWDB; approximately 150 HD and 42 VD endpoints are described by Interface2 v7. | HWDB component/install records
{Afe} | AFE block | 0..4 | DAPHNE hardware/protobuf
{Channel} | Readout channel | HD: 0..39; VD: 0..31. Do not instantiate absent VD channels. | HWDB channel map + detector profile
{Fan} | Fan | 0..1 | FpgaRegDict fanReadSpeed_0/1
{Interface} | Linux network interface | Baseline eth0; add only deployed/managed interfaces. | deployment/network inventory
{DataLink} | 10 GbE readout lane | Provisioned subset of 0..3 | HWDB/link configuration
{Sfp} | SFP role | timing, ccm, data0..data3 as physically provisioned | HWDB + schematic
{Bus} | I2C bus | Expected Linux buses 1 and 2; resolve stable role rather than assuming the numeric bus forever. | device tree + board service
{Address} | I2C address | Allow-listed expected devices plus discovered registered devices. | status snapshot
{Device} | SPI device | current-monitor /dev/spidev3.0 baseline | device tree + preflight
{Rail} | PMBus rail | 3VD3@0x12, 2VA1@0x16, 3VA6@0x32, 1VD8@0x36 on the PL I2C bus | Daphne I2C driver / Rust collector
{Sensor} | Temperature sensor | Linux thermal zones, hwmon sensors, carrier MCP9808, PMBus regulator temperatures, and DOM-capable SFPs after allow-list review. | status collector + hardware inventory
{Service} | Required service | firmware.service, clockchip.service, endpoint.service, hermes.service, daphne.service, daphne-boot-ok.service; remoteproc/RPU entries | PetaLinux service chain

## I2C Device Inventory
Bus role | Current bus | Address | Device/function | Notes
PS/SOM | dynamic | 0x50 | K26 SOM identity EEPROM | Identity/enrollment; read-only
PS/carrier | dynamic | 0x18 expected; verify | MCP9808 carrier temperature sensor | Do not freeze address until schematic/runtime validation
PL monitor | 1 | 0x10 | ADS7138: 3V3PDS, 1V8PDS, VBIAS0..4 | Current C++ cached monitor
PL monitor | 1 | 0x17 | ADS7138: 1V8A, 3V3A, -5VA | Current C++ cached monitor
PL service | 2 | 0x12 | PJT004A0X43 3VD3 regulator | PMBus
PL service | 2 | 0x16 | PJT004A0X43 2VA1 regulator | PMBus
PL service | 2 | 0x32 | PJT004A0X43 3VA6 regulator | PMBus
PL service | 2 | 0x36 | PJT004A0X43 1VD8 regulator | PMBus
PL service | 2 | 0x70 | Timing clock generator | Sanity read 0xE6 == 0x06 in current preflight
PL service | 2 | 0x71 | Mezzanine I2C mux/expander | Selects AFE0..4 downstream branches
PL service | 2 | 0x72 | SFP I2C/GPIO expander | Identity/DOM and sideband design surface
HD mezz downstream | behind 0x71 | 0x40 | INA232 3.3 V monitor | Per AFE branch
HD mezz downstream | behind 0x71 | 0x41 | TCA9536 power GPIO | Per AFE branch
HD mezz downstream | behind 0x71 | 0x42 | INA232 5 V monitor | Per AFE branch

## Raw Register Map
Register pattern | Logical/physical address | Fields | Bits | Access | Production treatment
afeGlobalControl | 0x00000000 | RESET | 0 | R/W | Canonical AFE.Global.ResetAsserted; DAQ configuration reset
afeGlobalControl | 0x00000000 | POWERSTATE | 1 | R/W | Canonical AFE.Global.PowerState; SC-owned energization
afeGlobalControl | 0x00000000 | BUSY/BUSY_AFE0/BUSY_AFE12/BUSY_AFE34 | 2..4 | R | Publish typed busy states
afeControl_{Afe} | 0x00000004 + 0xC*AFE | DATA | 0..23 | R/W | Raw AFE register transport; expert only
afeDacTrim_{Afe} | 0x00000008 + 0xC*AFE | DATA | 0..31 | R/W | Use typed per-channel trim nodes
afeDacOffset_{Afe} | 0x0000000C + 0xC*AFE | DATA | 0..31 | R/W | Use typed per-channel offset nodes
dacGainBiasControl | 0x0C000000 | BUSY, GO | 0..1 | R/W | Publish busy; hide raw GO
dacGainBiasU50/U53/U5 | 0x0C000004..0x0C00000C | DATA, BUFFER, GAIN, CHANNEL | 0..15 | R/W | Raw DAC control; expert only
timestamp{0..3} | 0x1002D000..0x10030000 | VALUE | 0..15 | R | Assemble typed Timing.Timestamp
endpointClockControl | 0x04000000 (phys 0x84000000) | SOFT_RESET, MMCM_RESET, CLOCK_SOURCE | 0..2 | R/W | Typed DAQ/DTS timing configuration and reset
endpointClockStatus | 0x04000004 (phys 0x84000004) | MMCM0_LOCKED, MMCM1_LOCKED | 0..1 | R | Current status API
endpointControl | 0x04000008 | ADDRESS, RESET | 0..16 | R/W | Typed DAQ/DTS endpoint address and reset
endpointStatus | 0x0400000C (phys 0x8400000C) | FSM_STATUS, TIMESTAMP_OK | 0..4 | R | Current status/preflight
frontendControl | 0x08000000 | DELAYCTRL_RESET, SERDES_RESET, DELAY_EN_VTC | 0..2 | R/W | Expert/DAQ alignment only
frontendStatus | 0x08000004 | DELAYCTRL_READY | 0 | R | Canonical alignment readiness
frontendTrigger | 0x08000008 | GO | 0 | W | Software trigger; DAQ/expert
frontendDelay_{Afe} | 0x0800000C + 4*AFE | DELAY | 0..8 | R/W | Canonical AFE.AlignmentDelay
frontendBitslip_{Afe} | 0x08000020 + 4*AFE | BITSLIP | 0..3 | R/W | Canonical AFE.AlignmentBitslip
spyTriggerControl | 0x08000034 (phys 0x88000034) | SOURCE_SELECTOR, INHIBIT | [1:0], 2 | R/W | Publish typed readbacks. Guard writes as one atomic DAQ/expert operation; reset 0x3. Source 0 software, 1 external, 2 timing/ad-hoc, 3 legacy OR/all.
spyBuffer_{Afe}_{LocalChannel} | 0x10000000 + 0x9000*AFE + 0x1000*channel | DATA/DATAL/DATAH | 0..31 | R | Diagnostic payload, not DCS time-series data
tenGigabitSender | 0x18000000 | DATA | 0..3 | R/W | Publish decoded readout-lane state; raw write expert only
fanControl | 0x14000000 | FAN_CTRL | 0..7 | R/W | Local automatic control / maintenance override
fanReadSpeed_0/1 | 0x14000004/08 | SPEED | 0..11 | R | Canonical fan tach/raw and validated RPM
biasEnable | 0x1400000C | ENABLE | 0 | R/W | SC-owned bias energization; DAQ owns the numerical target
muxEnable | 0x14000010 | ENABLE0, ENABLE1 | 0..1 | R/W | Internal driver state; do not expose raw write
muxAddress | 0x14000014 | ADDRESS | 0..1 | R/W | Internal driver state; do not expose raw write
LED | 0x14000018 | LED | 0..5 | R/W | Commissioning only
GIT | 0x1400001C | GIT | 0..27 | R | Canonical Firmware.GitCommit
triggerEnableLow/High | 0x14000020/24 (phys 0x94000020/24) | DATA | 40-bit mask | R/W | Canonical per-channel TriggerEnabled
idLink/idSlot/idCrate/idDetector/idVersion | 0x14000028..38 | ID | field-specific | R/W | Identity readback; writes only via deployment/configure
adHocTriggerCommand | 0x1400003C | VALUE | 0..7 | W | DAQ/expert only
selfTriggerFullConfigLow/High | 0x14000040/44 | VALUE | 64-bit | R/W | Typed run configuration preferred
matchingTriggerTemplate_{0..15} | 0x14000048 + 4*i | VALUE | 0..27 | R/W | DAQ calibration/configuration; version and read back
triggerCounters_{Channel} | 0xA0010000 + 0x20*channel | threshold, record, busy, full | documented offsets | R | Current MT2_READ_TRIGGER_COUNTERS
selfTriggerStuff | 0x94000000 region | config, delay, filter select, compensator, inverter | documented offsets | R/W | Typed run configuration preferred
DTS command decoder | control bus 0x60..0x67 | tcmd_0..tcmd_7 | 8-bit each | R/W | DAQ/DTS-owned command-to-output mapping; EDMS 2395364 v2
DTS endpoint command | control bus 0x70 | disable/enable TX, set address, set coarse delay, resync, reset | bits 0..5 | W | DAQ/DTS-owned endpoint control; timing master has bus precedence
DTS endpoint status | control bus 0x71 | FSM status, phase-update complete, transmitter enabled | bits 0..5 | R | Publish to DAQ and SC; EDMS 2395364 v2
DTS endpoint deskew | control bus 0x72..0x73 | coarse delay, fine clock phase | bits 3..0, 15..4 | R/W | DAQ/DTS-owned deskew configuration
DTS endpoint address | control bus 0x74..0x75 | unique endpoint address | 16-bit | R/W | DAQ/DTS-owned endpoint identity
DTS protocol errors | control bus 0x76 | soft protocol-error count; write clears | 8-bit | R/W | DAQ/DTS owns clear; DAQ and SC monitor
DTS timestamp check | control bus 0x77 | timestamp set/check FIFO value | 8-bit transaction | R/W | DAQ/DTS-owned timing operation
DTS endpoint version | control bus 0x78..0x7B | endpoint version | 32-bit | R | Publish to DAQ and SC

## Protocol Surface
Message/operation family | Class | Contents | Normal authority | Status | Notes
MT2_READ_TEST_REG | Read | Service heartbeat/test value | Read-only fan-out | CURRENT OPC-UA POC | TestRegResponse
MT2_READ_GENERAL_INFO | Read | VBIAS0..4, selected rails, temperature field | Read-only fan-out | CURRENT OPC-UA POC | GeneralInfo; current C++ handler does not populate temperature
MT2_READ_SLOW_CONTROL_STATUS | Read | Firmware, RPU, I2C, clocks, temperatures, rails, services, errors | Read-only fan-out | CURRENT PROTOBUF | Rust additive status API
MT2_READ_TRIGGER_COUNTERS | Read | Per-channel threshold/record/busy/full | Read-only; DAQ context | CURRENT PROTOBUF | 40 channels
READ_AFE_REG/VGAIN/BIAS/TRIM/OFFSET/VBIAS | Read | AFE and channel configuration readbacks | Read-only; serialized | CURRENT PROTOBUF | Low-level protobuf
READ_CURRENT_MONITOR | Read | Per-channel current monitor raw value | Read-only; serialized | CURRENT PROTOBUF | Low-level protobuf/SPI
READ_BIAS_VOLTAGE_MONITOR | Read | Per-AFE bias monitor | Read-only; serialized | CURRENT PROTOBUF | Low-level protobuf/I2C cache
READ_HDMEZZ_STATUS/CONFIG | Read | Per-AFE HD-mezz power, measurements, alerts, calibration config | Read-only; cached | CURRENT PROTOBUF | Low-level protobuf
CONFIGURE_FE | Backend write | Complete AFE/channel/run configuration | DAQ | CURRENT PROTOBUF | Mapped from DAQ-role OPC-UA ConfigureRun/ConfigureFrontend
CONFIGURE_CLKS | Backend write | Endpoint address, source and bounded resets | DAQ/DTS | CURRENT PROTOBUF | Mapped from DAQ/DTS-role OPC-UA ConfigureClocks
WRITE_AFE_BIAS_CONTROLLED_SET | Backend write | Closed-loop numerical bias target | DAQ | CURRENT PROTOBUF | Mapped from DAQ-role OPC-UA SetControlledBias; requires SC/DPS permit
SET_AFE_POWERSTATE | Backend write | AFE global power enable/state | Slow Controls | CURRENT PROTOBUF | Mapped from SC-role OPC-UA SetAfePowerState
SET_HDMEZZ_POWER_STATES | Write | Paired 5 V/3.3 V request | Slow Controls | CURRENT PROTOBUF | Never split into independent writes
CONFIGURE_HDMEZZ_BLOCK | Write | Shunt/current/protection configuration | Slow Controls with PDS expert role | CURRENT PROTOBUF | SC maintenance only
ALIGN_AFE | Write | Delay/bitslip alignment | DAQ/expert | CURRENT PROTOBUF | Not while running
DUMP_SPYBUFFER | Diagnostic | Waveform data transfer | DAQ/expert | EXCLUDED FROM ROUTINE DCS | Not a DCS historian path
DO_SOFTWARE_TRIGGER | Diagnostic | Ad hoc trigger | DAQ/expert | EXCLUDED FROM ROUTINE DCS | Not routine DCS
SET_SPY_TRIGGER_CONTROL | Backend write | Atomic spy-buffer source selector and inhibit | DAQ | REQUIRED PRODUCTION | DAQ-role OPC-UA method; backend adapter is not yet available in the reviewed Protobuf. Inhibit, wait >=5 acquisition clocks, select, wait >=5 clocks, then release inhibit when requested.
WRITE_MULTIPLE_REGISTER | Raw write | Arbitrary address/value list | PDS expert | EXCLUDED FROM ROUTINE DCS | No production Ignition exposure

## OPC-UA Control Policy
Operation | NodeId pattern | Control owner | Allowed roles | Implementation status | Backend mapping | Executable | Input contract | Output contract
RequestSafeState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.RequestSafeState | Slow Controls / DPS | sc,dps,expert | REQUIRED PRODUCTION |  | false | JSON request object | JSON response object
AcknowledgeRecovery | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.AcknowledgeRecovery | Slow Controls | sc,expert | REQUIRED PRODUCTION |  | false | JSON request object | JSON response object
ConfigureRun | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureRun | DAQ | daq,expert | CURRENT PROTOBUF | MT2_CONFIGURE_FE_REQ | true | JSON request object | JSON response object
ConfigureClocks | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureClocks | DAQ / DTS | daq,dts,expert | CURRENT PROTOBUF | MT2_CONFIGURE_CLKS_REQ | true | JSON request object | JSON response object
ConfigureFrontend | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureFrontend | DAQ | daq,expert | CURRENT PROTOBUF | MT2_CONFIGURE_FE_REQ | true | JSON request object | JSON response object
SetControlledBias | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetControlledBias | DAQ | daq,expert | CURRENT PROTOBUF | MT2_WRITE_VBIAS_CONTROL_REQ | true | JSON request object | JSON response object
SetAfePowerState | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetAfePowerState | Slow Controls | sc,expert | CURRENT PROTOBUF | MT2_SET_AFE_POWERSTATE_REQ | true | JSON request object | JSON response object
ResetAfe | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ResetAfe | DAQ | daq,expert | CURRENT PROTOBUF | MT2_DO_AFE_RESET_REQ | true | JSON request object | JSON response object
AlignAfes | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.AlignAfes | DAQ | daq,expert | CURRENT PROTOBUF | MT2_ALIGN_AFE_REQ | true | JSON request object | JSON response object
SetHdMezzPowerStates | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetHdMezzPowerStates | Slow Controls | sc,expert | CURRENT PROTOBUF |  | false | JSON request object | JSON response object
ConfigureHdMezzProtection | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ConfigureHdMezzProtection | Slow Controls | sc,expert | EXCLUDED FROM ROUTINE DCS |  | false | JSON request object | JSON response object
ClearHdMezzAlert | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ClearHdMezzAlert | Slow Controls | sc,expert | CURRENT PROTOBUF |  | false | JSON request object | JSON response object
SetFanOverride | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetFanOverride | Slow Controls | sc,expert | EXCLUDED FROM ROUTINE DCS |  | false | JSON request object | JSON response object
ScrapEndpoint | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.ScrapEndpoint | DAQ | daq,expert | CURRENT PROTOBUF |  | false | JSON request object | JSON response object
SoftwareTrigger | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SoftwareTrigger | DAQ | daq,expert | CURRENT PROTOBUF | MT2_DO_SOFTWARE_TRIGGER_REQ | true | JSON request object | JSON response object
SetSpyTriggerControl | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.SetSpyTriggerControl | DAQ | daq,expert | REQUIRED PRODUCTION |  | false | JSON request object | JSON response object
DumpSpyBuffers | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.DumpSpyBuffers | DAQ | daq,expert | EXCLUDED FROM ROUTINE DCS | MT2_DUMP_SPYBUFFER_REQ | true | JSON request object | JSON response object
WriteAfeRegister | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.WriteAfeRegister | PDS expert (maintenance only) | expert | EXCLUDED FROM ROUTINE DCS | MT2_WRITE_AFE_REG_REQ | true | JSON request object | JSON response object
WriteMultipleFpgaRegisters | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.WriteMultipleFpgaRegisters | PDS expert (maintenance only) | expert | EXCLUDED FROM ROUTINE DCS |  | false | JSON request object | JSON response object
FirmwareUpdate | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.FirmwareUpdate | Slow Controls | sc,expert | EXCLUDED FROM ROUTINE DCS |  | false | JSON request object | JSON response object
RestartBoardService | nsu=urn:dune:pds:daphne;s=DAPHNE.Boards.{BoardId}.Operations.Methods.RestartBoardService | Slow Controls | sc,expert | REQUIRED PRODUCTION |  | false | JSON request object | JSON response object

## Operations Constraints
External system/run mode | Required input/state | DAPHNE/PDS guard | Owner | Failure/overtake action | Return condition
Cryogenics | LAr level valid/above approved threshold; cryogenic state stable | SC/DPS permit to bias/operate relevant PDS loads | Cryogenics + DPS | Remove permit; SC ramps/disables affected bias or optical power through the safe-state path | Permit valid, condition stable for agreed dwell, SC acknowledgement, then fresh DAQ configuration through OPC-UA
High voltage | Cathode/HV state, ramping, trip/discharge, approved operating mode | VD cathode PoF/SiPM bias and calibration permissions | HV + DPS | Block transitions or force the approved SC safe state for affected cathode channels | HV stable/approved, DPS permit restored, SC recovery complete, then fresh DAQ configuration through OPC-UA
Rack cooling | Airflow/fan/rack-temperature and cooling availability | DAPHNE run-ready and local thermal control | Slow Controls/DPS | Alarm; remove run-ready; inhibit or shut down per reviewed thermal analysis | Cooling restored, temperatures in range for dwell, acknowledgement/reconfigure
PDU/UPS/power | Supply healthy, breaker/PDU/UPS state, external LV stable | DAPHNE power and DAQ-configuration permit | Slow Controls/DPS | Block DAQ configuration or request SC safe state; hard inhibit where implemented | Stable power and permits; fresh DAQ configuration through OPC-UA
Physics run | Run Control RUNNING with accepted configuration hash, target scope, and valid SC/DPS permits | All run-affecting AFE/timing/trigger writes | DAQ | Reject non-DAQ writes to data-taking fields; SC monitoring, equipment control and protection continue without tuning DAQ-owned settings | Run stopped or DAQ change accepted through the approved state machine
Noise run | RunType=Noise, versioned noise configuration, target scope and valid SC/DPS permits | Trigger target, numerical bias target and AFE configuration | DAQ | SC does not tune DAQ-owned values; retain complete monitoring and label history with run/config revision | Fresh versioned DAQ physics/calibration profile
Calibration run | RunType=Calibration; LCM, timing, DAPHNE and external permits ready | LCM pulse and DAPHNE capture/configuration sequence | DAQ orchestrates through its OPC-UA role; SC supplies equipment readiness and permits | DPS/SC veto aborts or safe-states; synchronous DTS command performs the timed action | Calibration completes/aborts, SC safe restoration is verified, and DAQ freshly applies the physics profile through OPC-UA
Timing | DTS lock, accepted endpoint FSM, timestamp valid | Timing configuration and data-taking readiness | DAQ/DTS | Block DAQ configuration/data-ready; attempt bounded firmware auto-relock without perturbing locked endpoints; SC monitors | Timing usable and fresh DAQ configuration through OPC-UA
External activity compatibility | Fresh authoritative active/requested states, detector-region scope, and an approved site-specific compatibility-matrix revision | ExternalActivityPermit for PDS/SiPM-bias/PoF activation; source states remain owned by the external systems | Slow Controls/DPS admission gate; subsystem owners supply state and cause-and-effect | Reject a newly requested activation when an approved pair is incompatible. Block new activation on stale or unresolved inputs; the response for already-active equipment requires an approved cause-and-effect rule and shall not be inferred from a color cell. | Conflict cleared, inputs fresh, permits restored, approved acknowledgement/dwell completed, and fresh DAQ configuration applied when prior configuration was invalidated
Maintenance | SC maintenance lock and physical access/work permit | Expert diagnostics, raw registers, firmware updates | Slow Controls + PDS expert | Block routine DAQ and SC OPC-UA methods; preserve audit and safe state | Maintenance released, checks passed, SC/DPS permits valid, then DAQ configuration through OPC-UA

## NP02 Compatibility
Site | Active source label | Active working meaning | Requested source label | Requested working meaning | Source classification | Proposed review treatment (not source text) | DAPHNE/PDS relevance or conflict | Missing production evidence | Source reference
NP02 | PoF | Power-over-Fiber system | LEDs | LED system; the PDF does not define whether this means camera illumination, calibration LEDs, or both | Incompatible | Admission evidence: deny the newly requested activation while the peer is active; define priority, timeout, already-active response, and recovery before implementation. | Define the LED class and affected PoF/PDS region before implementing an inhibit. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PoF | Power-over-Fiber system | PMTs | NP02 photomultiplier-tube system | Probably compatible | Unresolved: treat as not approved for automatic admission until the owners validate and reclassify the pair in a controlled matrix. | Unresolved NP02 pairing; no automatic permission may be inferred. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PoF | Power-over-Fiber system | PDs | NP02 photon-detector system | Compatible | No pairwise exclusion indicated by this source; all other permits, state guards, and safety procedures remain applicable. | Expected PDS dependency; still requires approved PoF startup, health, and region mapping. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PoF | Power-over-Fiber system | IoL | Ionisation-laser system | Incompatible | Admission evidence: deny the newly requested activation while the peer is active; define priority, timeout, already-active response, and recovery before implementation. | Candidate PoF admission inhibit during ionisation-laser operation; priority and already-active response are absent from the source. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PoF | Power-over-Fiber system | PrM | Purity-monitor system | Probably compatible | Unresolved: treat as not approved for automatic admission until the owners validate and reclassify the pair in a controlled matrix. | Unresolved; the SC-PDS draft discusses PrM versus PDS bias but does not specify this PoF pairing. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | LEDs | LED system; the PDF does not define whether this means camera illumination, calibration LEDs, or both | PMTs | NP02 photomultiplier-tube system | Incompatible | Admission evidence: deny the newly requested activation while the peer is active; define priority, timeout, already-active response, and recovery before implementation. | Define whether this is camera illumination rather than detector calibration light. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | LEDs | LED system; the PDF does not define whether this means camera illumination, calibration LEDs, or both | PDs | NP02 photon-detector system | Incompatible | Admission evidence: deny the newly requested activation while the peer is active; define priority, timeout, already-active response, and recovery before implementation. | Consistent with the intended camera-LED/PDS exclusion in SC-PDS draft v1, whose second sentence incorrectly names the Purity Monitor. Do not apply this row to PDS calibration LEDs without scope approval. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | LEDs | LED system; the PDF does not define whether this means camera illumination, calibration LEDs, or both | IoL | Ionisation-laser system | Compatible | No pairwise exclusion indicated by this source; all other permits, state guards, and safety procedures remain applicable. | No pairwise exclusion is shown; normal safety and operating guards still apply. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | LEDs | LED system; the PDF does not define whether this means camera illumination, calibration LEDs, or both | PrM | Purity-monitor system | Compatible | No pairwise exclusion indicated by this source; all other permits, state guards, and safety procedures remain applicable. | No pairwise exclusion is shown; normal safety and operating guards still apply. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PMTs | NP02 photomultiplier-tube system | PDs | NP02 photon-detector system | Compatible | No pairwise exclusion indicated by this source; all other permits, state guards, and safety procedures remain applicable. | NP02 evidence permits coexistence; it is not an FD operating requirement. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PMTs | NP02 photomultiplier-tube system | IoL | Ionisation-laser system | Incompatible | Admission evidence: deny the newly requested activation while the peer is active; define priority, timeout, already-active response, and recovery before implementation. | Candidate NP02 admission exclusion; signal, scope, and enforcement path are not defined. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PMTs | NP02 photomultiplier-tube system | PrM | Purity-monitor system | Probably compatible | Unresolved: treat as not approved for automatic admission until the owners validate and reclassify the pair in a controlled matrix. | Unresolved NP02 pairing; no automatic permission may be inferred. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PDs | NP02 photon-detector system | IoL | Ionisation-laser system | Incompatible | Admission evidence: deny the newly requested activation while the peer is active; define priority, timeout, already-active response, and recovery before implementation. | Candidate missing SC-PDS external constraint: ionisation-laser operation versus PDS activation/bias. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | PDs | NP02 photon-detector system | PrM | Purity-monitor system | Probably compatible | Unresolved: treat as not approved for automatic admission until the owners validate and reclassify the pair in a controlled matrix. | CONFLICT: SC-PDS ICD draft v1 requires PDS SiPM bias off while the Purity Monitor is active. Resolve before either artifact is approved. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9
NP02 | IoL | Ionisation-laser system | PrM | Purity-monitor system | Compatible | No pairwise exclusion indicated by this source; all other permits, state guards, and safety procedures remain applicable. | No pairwise exclusion is shown; normal safety and operating guards still apply. | Source supplies no author, document/revision ID, variable names, scope mapping, timing, priority, failure action, recovery, or approval record. | NP02interlockMatrix.pdf; PDF created 2025-03-10; SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9

## Source Audit
Source | Revision | Files/area reviewed | Use and limitation
daphneZMQ | b18075ed3e7d | srcs/protobuf/daphneV3_*; FpgaRegDict.cpp; defines.hpp; monitoring/handlers; HD-mezz docs | Current C++ control/monitor and register surface; working tree has unrelated emulator changes
daphnemodules | develop edf75a0c5f37 (2026-04-02; upstream current when checked 2026-08-13) | daphne_control_high/low.proto; opmon schema; DaphneV2ControllerModule configuration sequence | Defines the DAQ configuration fields and sequence used to design the OPC-UA methods; DAQ should use the gateway rather than maintain the DAPHNE backend Protobuf
daphne-sc | a3df24e0e318 + shared-gateway working tree | daphne_sc.proto; Rust status collector/preflight/RPU; OPC-UA bridge | Current NP04 bridge is a read-only proof of concept; this revision adds workbook-driven roles and mapped configuration methods
daphne-firmware | 42534e50f145 + reviewed working tree | Fpga service chain; endpoint contract; Memory_Map.md; docs/modules/spy-buffer.md; fe_axi.vhd; board identity/network deployment | Firmware, services, identity, network, and gateware requirements. The spy-trigger selector/inhibit at 0x88000034 is a working-tree interface not yet exposed by Protobuf/ZMQ.
DTS Endpoint Integration Guide | EDMS 2580777 file v1; document revision v2 dated 2025-10-17 | Physical layer, TX_DISABLE, LOS, SFP I2C/DOM, endpoint integration and status readiness | Authoritative timing-interface guide, including the endpoint SFP specification: TX 1550 nm / RX 1490 nm
DTS Endpoint Firmware Interface | EDMS 2395364 file v2; document revision v6 dated 2025-10-17 | Endpoint generics, ports, startup FSM, control-bus arbitration, command decoder and register map | Defines DAQ/DTS-owned address, deskew, command mapping, resync/reset and transmitter request plus monitored status/version/errors
DUNE Timing Interface Definition | EDMS 2209895 v1 (2018) | Power-on initialization and continuous monitoring responsibilities | Historical interface source used only for the principle that timing monitored quantities are sent to SC; superseded optical wavelength text is not used
daphne_slow_controls_plan | 31c289179d5d | main.tex; VST variable registry; SC-DPS plan | Production variable families, alarms/history, and external constraints; DAQ and SC use a shared role-aware OPC-UA gateway
Interface2 | EDMS 2088726 proposed v8.0 | Original-format LaTeX revision in the Interface2 repository | DAQ-PDS physical/protocol boundary, shared OPC-UA gateway, ownership and scaling
DCS_Intake_Template.xlsx | local supplied copy | sc_interface/gizmo/sources | Slow Controls intake column layout only; example values are cleared
CERN EDMS navigator collection supplied by requester | checked 2026-08-13 | https://edms.cern.ch/ui/#!master/navigator/document?P:100905576:101799625:subDocs | The timing documents used for the ownership allocation are enumerated explicitly above; remaining installed values and thresholds still require controlled sign-off
NP02 activity context | Vertical Drift TB, 2025-02-26; Filippo Resnati | https://indico.cern.ch/event/1522253/contributions/6404491/attachments/3026376/5341354/TB_Feb25.pdf | Official context identifies PDS commissioning with PoF laser safety approval and installation of the ionisation laser. It supports working subsystem meanings only and does not approve the supplied compatibility classifications.
NP02 compatibility matrix | PDF created 2025-03-10; supplied 2026-08-17 | NP02interlockMatrix.pdf; PoF, LEDs, PMTs, PDs, IoL, and PrM pairwise activation classifications | Historical NP02 admission evidence only; not an FD cause-and-effect specification. The source lacks author, document/revision ID, definitions, scope, signals, timing, priority, failure/recovery behavior, and approval. SHA-256 0e4027da22fb88182626a7baff22cf632d1e72a3183d93f39b743ca2c109cfa9